## WiredTestReport
ID | Title | Document | Type | Relationship | Item
UTICA-25500 | Overall, 12 of 13 (92.3%) participants reported the battery life to be acceptable. | Wired Test Report | User Validation Case | refines | UTICA-9719
UTICA-25511 | Yes, participants were able to successfully insert and remove the devices from their ears. The average participant ratings for ease of insertion and removal were 6.23 and 6.85, respectively, on a 7-pt scale, with 1 representing "very difficult" and 7 representing "very easy." | Wired Test Report | User Validation Case | refines | UTICA-11458
UTICA-25509 | Yes, participants were able to successfully insert and remove a battery in the device. The average participant ratings for the function and ease of use of the battery door for insertion and removal were 5.85 and 5.00, respectively, on a 7-pt scale, with 1 representing "very difficult" and 7 representing "very easy." | Wired Test Report | User Validation Case | refines | UTICA-19527
UTICA-25517 | Yes, participants were able to successfully insert and remove the hearing aid devices from their hearing aid cases. The average participant ratings for ease of use and transport for the hearing aid case were 6.85 and 6.77, respectively, on a 7-pt scale, with 1 representing "very difficult" and 7 representing "very easy." | Wired Test Report | User Validation Case | refines | UTICA-11459
UTICA-25515 | Yes, participants were able to successfully activate the user controls. The average participant rating for the function and/or ease of use of the user controls was 6.33, on a 7-pt scale, with 1 representing "very difficult" and 7 representing "very easy." | Wired Test Report | User Validation Case | refines | UTICA-9707
UTICA-25529 | Yes, participants were successfully able to clean the hearing aids. Overall, the average participant rating for the function and ease of use of cleaning the devices was 5.67 on a 7-pt scale, with 1 representing "very difficult" and 7 representing "very easy." | Wired Test Report | User Validation Case | refines | UTICA-9709
UTICA-25525 | Yes. The average participant rating for the physical comfort of the devices was 6.46, on a 7-pt scale, with 1 representing "very poor" and 7 representing "very good." | Wired Test Report | User Validation Case | refines | UTICA-9710
UTICA-25533 | The average SII as determined via real ear measures for REM Target Match was 68.91 and the average SII as determined via real ear measures for manual (traditional) real ear measurements based on Best Fit was 69.09. The average SII for REM Target Match is as good as the average SII for manual (traditional) measurements based on Best Fit (+/- 3 dB from targets). | Wired Test Report | User Validation Case | refines | UTICA-9852
UTICA-26344 | The average time to complete REM Target Match procedures with a Verifit 2 was 3 minutes and 15 seconds. This was less than the average time (7 minutes 8 seconds) to complete manual (traditional) real ear measurements with Verifit 2. | Wired Test Report | User Validation Case | refines | UTICA-6792
UTICA-25599 | Yes, the average participant responses on the DOSO indicate benefit with Utica wired study test devices over own personal devices. | Wired Test Report | User Validation Case | refines | UTICA-16344
UTICA-25521 | The average participant rating for overall satisfaction for the hearing aids was 7.92 on an 11-pt scale, with 0 representing "not at all satisfied" and 10 representing "extremely satisfied." The average participant rating for likelihood to recommend was 7.85, on an 11-pt scale, with 0 representing "not at all likely" and 10 representing "extremely likely." Overall, these ratings appear to be similar to previous products, but the statistical analysis could not be completed due to scaling differences and the limited size of the population. | Wired Test Report | User Validation Case | refines | UTICA-16345
UTICA-25521 | The average participant rating for overall satisfaction for the hearing aids was 7.92 on an 11-pt scale, with 0 representing "not at all satisfied" and 10 representing "extremely satisfied." The average participant rating for likelihood to recommend was 7.85, on an 11-pt scale, with 0 representing "not at all likely" and 10 representing "extremely likely." Overall, these ratings appear to be similar to previous products, but the statistical analysis could not be completed due to scaling differences and the limited size of the population. | Wired Test Report | User Validation Case | is refined by | UTICA-25551
UTICA-25553 | No adverse events were reported during the design validation. | Wired Test Report | User Validation Case | refines | UTICA-9703
UTICA-25551 | Yes, based on the Utica wired IIC and CIC design validation results, the products are performing as designed for the intended population. | Wired Test Report | User Validation Case | refines | UTICA-9702
UTICA-25551 | Yes, based on the Utica wired IIC and CIC design validation results, the products are performing as designed for the intended population. | Wired Test Report | User Validation Case | refines | UTICA-25521

## Source
ID | Title | Type | Linked Work Items
UTICA-25550 | Yes, when averaged across all participants, there was a significant difference between REM Target Match aided mean Speech Intelligibility Index (SII) values compared to the unaided mean SII values obtained via real-ear measurements. | User Validation Case | has parent: UTICA-25552 - 9.1.1
UTICA-25558 | On a 7-points scale, the overall average participant rating for sound quality was 6.46 in quiet and 4.92 in noisy environments. Overall, 12 of 13 (92.3%) participants reported sound quality in quiet to be "good" or "very good," and 6 of 13 (46.2%) participants reported sound quality in noise to be "good" or "very good." Also, on a 7-point scale, the average participant rating for speech understanding was 6.54 in quiet and 4.92 in noisy environments. Overall, 12 of 13 (92.3%) participants rated speech understanding in quiet to be "good" or "very good," and 8 of 13 (61.5%) participants rated speech understanding in noise to be "good" or "very good." | User Validation Case | has parent: UTICA-25560 - 9.1.2
UTICA-25502 | Yes, 7 of 13 (53.8%) participants reported unexpected or objectionable artifacts while in the field. The average participant rating for overall frequency was 3.0 (sometimes) on a scale of 1-5, with 1 representing "never" and 5 representing "always." The average participant rating for the overall severity was 3.9 (between moderate and irritant), with 1 representing "unusable" and 5 representing "not bothersome." | User Validation Case | has parent: UTICA-25503 - 9.1.3
UTICA-25500 | Overall, 12 of 13 (92.3%) participants reported the battery life to be acceptable. | User Validation Case | refines: UTICA-9719, has parent: UTICA-25501 - 9.1.4
UTICA-25511 | Yes, participants were able to successfully insert and remove the devices from their ears. The average participant ratings for ease of insertion and removal were 6.23 and 6.85, respectively, on a 7-pt scale, with 1 representing "very difficult" and 7 representing "very easy." | User Validation Case | refines: UTICA-11458, has parent: UTICA-25512 - 9.2.1
UTICA-25509 | Yes, participants were able to successfully insert and remove a battery in the device. The average participant ratings for the function and ease of use of the battery door for insertion and removal were 5.85 and 5.00, respectively, on a 7-pt scale, with 1 representing "very difficult" and 7 representing "very easy." | User Validation Case | refines: UTICA-19527, has parent: UTICA-25510 - 9.2.2
UTICA-25517 | Yes, participants were able to successfully insert and remove the hearing aid devices from their hearing aid cases. The average participant ratings for ease of use and transport for the hearing aid case were 6.85 and 6.77, respectively, on a 7-pt scale, with 1 representing "very difficult" and 7 representing "very easy." | User Validation Case | refines: UTICA-11459, has parent: UTICA-25518 - 9.2.3
UTICA-25515 | Yes, participants were able to successfully activate the user controls. The average participant rating for the function and/or ease of use of the user controls was 6.33, on a 7-pt scale, with 1 representing "very difficult" and 7 representing "very easy." | User Validation Case | refines: UTICA-9707, has parent: UTICA-25516 - 9.2.4
UTICA-25529 | Yes, participants were successfully able to clean the hearing aids. Overall, the average participant rating for the function and ease of use of cleaning the devices was 5.67 on a 7-pt scale, with 1 representing "very difficult" and 7 representing "very easy." | User Validation Case | refines: UTICA-9709, has parent: UTICA-25530 - 9.2.5
UTICA-25525 | Yes. The average participant rating for the physical comfort of the devices was 6.46, on a 7-pt scale, with 1 representing "very poor" and 7 representing "very good." | User Validation Case | refines: UTICA-9710, has parent: UTICA-25598 - 9.2.6
UTICA-25533 | The average SII as determined via real ear measures for REM Target Match was 68.91 and the average SII as determined via real ear measures for manual (traditional) real ear measurements based on Best Fit was 69.09. The average SII for REM Target Match is as good as the average SII for manual (traditional) measurements based on Best Fit (+/- 3 dB from targets). | User Validation Case | refines: UTICA-9852, has parent: UTICA-25535 - 9.3.1
UTICA-26344 | The average time to complete REM Target Match procedures with a Verifit 2 was 3 minutes and 15 seconds. This was less than the average time (7 minutes 8 seconds) to complete manual (traditional) real ear measurements with Verifit 2. | User Validation Case | refines: UTICA-6792, has parent: UTICA-26343 - 9.3.2
UTICA-25599 | Yes, the average participant responses on the DOSO indicate benefit with Utica wired study test devices over own personal devices. | User Validation Case | refines: UTICA-16344, has parent: UTICA-25523 - 9.4.1
UTICA-25521 | The average participant rating for overall satisfaction for the hearing aids was 7.92 on an 11-pt scale, with 0 representing "not at all satisfied" and 10 representing "extremely satisfied." The average participant rating for likelihood to recommend was 7.85, on an 11-pt scale, with 0 representing "not at all likely" and 10 representing "extremely likely." Overall, these ratings appear to be similar to previous products, but the statistical analysis could not be completed due to scaling differences and the limited size of the population. | User Validation Case | refines: UTICA-16345, has parent: UTICA-25522 - 9.4.2, is refined by: UTICA-25551 - 9.4.4-1
UTICA-25553 | No adverse events were reported during the design validation. | User Validation Case | refines: UTICA-9703, has parent: UTICA-25557 - 9.4.3
UTICA-25551 | Yes, based on the Utica wired IIC and CIC design validation results, the products are performing as designed for the intended population. | User Validation Case | refines: UTICA-9702, refines: UTICA-25521 - 9.4.2-1, has parent: UTICA-25555 - 9.4.4