## 900 MHz
Product | Matrix | Firmware Rev. | Battery Size | Nominal Current FT/SHA off (µA) | Nominal Current FT/SHA (µA) | CVSD Stream Current (µA) | G722 Stream Current (µA) | E2E Transmit Current (µA) | E2E Receive Current (µA) | Bidirectional Current (µA) | WiCross Xmt Current (µA) | WiCross Rcv Current (µA) | Expected Days Typ Accessory | Actual Days Typ Accessory | Expected Days Non Accessory | Actual Days Non Accessory | Serial Numbers | LiIon supply voltage | Vbat Supply Voltage | Notes | NOTE: * = Calculated
Kiev RIC 312 | 115/50 | 4.0.5.1 | 312 | 1450 |  | 2575 | 2975 | 3025 | 2800 | 3775 | 2825 | 2850 | NA |  | *5.2 |  | 15373216 / 15373186 |  | 1.2 | Typical Wireless Accessory
[13 empty] | NA | [5 empty] | 1.2
| 123/60 | [11 empty] | NA | [5 empty] | 1.2
[13 empty] | NA | [5 empty] | 1.2
| 130/70 | 4.0.5.1 | 312 | 1500 |  | 2625 | 3025 | 3050 | 2800 | 3750 | 2850 | 2875 | NA |  |  |  | 15373216 / 15373186 |  | 1.2
Kiev BTE 312 | 130/60 | 4.0.5.1 | 312 | 1400 |  | 2475 | 2875 | 3000 | 2725 | 3675 |  |  | NA |  |  |  | 15117925 / 15117926 |  | 1.2
[13 empty] | NA | [5 empty] | 1.2
Kiev ITE 13 | 115/50 | [11 empty] | NA | [5 empty] | 1.2
| 123/60 | [11 empty] | NA | [5 empty] | 1.2
| 130/70 | 4.0.5.1 | 13 | 1425 |  | 2625 | 3050 | 3075 | 2775 | 3800 | 2900 | 2850 | NA |  |  |  | 2911321707 / 2911321708 |  | 1.2
Kiev ITC 312 | 110/40 | 4.0.5.1 | 312 | 1300 |  | 2500 | 2925 | 3000 | 2700 | 3800 | 2900 | 2800 | NA |  |  |  | 2911321781 / 2911321782 |  | 1.2
| 115/50 | 4.0.5.1 | 312 | 1350 |  | 2550 | 3000 | 2950 | 2800 | 3800 | 2825 | 2900 | NA |  |  |  | 2911321806 / 2911321807 |  | 1.2
| 120/60 | 4.0.5.1 | 312 | 1350 |  | 2450 | 2850 | 2950 | 2700 | 3625 | 3050 | 2800 | NA |  |  |  | 2911321737 / 2911321738 |  | 1.2
[13 empty] | NA | [5 empty] | 1.2
Kiev CIC 10 | 110/40 | 4.0.5.1 | 10 | 1250 |  | 2425 | 2825 | 2800 | 2625 | 3500 | NA | NA | NA |  |  |  | 2911321723 / 2911321724 |  | 1.2
| 115/50 | [11 empty] | NA | [5 empty] | 1.2
SoundLens IIC | 110/40 | 4.3.1.1 | 10 | 1250 |  | 2425 | 2825 | 2875 | 2600 | 3525 | NA | NA | NA |  |  |  | 2911323833 / 2911323834 |  | 1.2
| 115/50 | [11 empty] | NA | [5 empty] | 1.2
London RIC 312 | 110/40 | [11 empty] | NA | [5 empty] | 1.2
| 115/50 | 4.3.2.1 | 312 | 1550 |  | 2700 | 3100 | 3225 | 2750 |  | NA | NA | NA |  |  |  | 2911323833 / 2911323834 |  | 1.2
| 115/50 | 4.3.2.3 | 312 | 1475 |  | 2575 | 3000 | 3125 | 2750 | 3725 | NA | NA | *6.7 |  |  |  | 15118002 / 41181290 |  | 1.2
| 120/60 | 4.10.0.2 | 312 | 1575 |  | 2825 |  |  |  | 3750 | [6 empty] | 160159070 |  | 1.2 | Angled receiver
Sapphire Lyon BTE 13 | 115/50 | [11 empty] | NA | [5 empty] | 1.2
| 123/60 | [11 empty] | NA | [5 empty] | 1.2
| 130/70 | 4.3.2.8 | 13 | 1600 |  | 2600 | 2970 | [5 empty] | NA |  |  |  | 2911321707 / 2911321708 |  | 1.2
| 130/70 | 4.3.2.8 | 13 | 1400 |  | 2725 | 3100 |  |  | 3750 |  |  | NA |  |  |  | 160499662 / 160499652 |  | 1.2
Muse BTE 312 | 130/60 | 4.3.2.10 | 312 | 1350 |  | 2525 | 2925 |  |  | 3675 |  |  | NA | [5 empty] | 1.2
|  | 4.3.2.10 | 312 | 1375 |  | 2525 | 2925 | 3150 | 2750 | 3750 | NA | NA | NA |  |  |  | 15613951 / 15613953 |  | 1.2
|  | 4.3.2.10* | 312 | 1400 |  | 2600 | NA | 3150 | 2700 | 3875 | NA | NA | NA |  |  |  | 151423058 / 151423078 |  | 1.2 | * EU device current
Muse BTE 13 | 130/60 | 4.3.2.10 | 13 | [9 empty] | NA | [5 empty] | 1.2
|  |  | 13 | [9 empty] | NA | [5 empty] | 1.2
Muse RIC 312 | 110/40 | [11 empty] | NA | [5 empty] | 1.2
| 115/50 | 4.3.2.10 | 312 | 1475 |  | 2650 | 3050 | 3175 | 2750 | 3800 |  |  | NA |  |  |  | 15373307 / 15373236 |  | 1.2
| 115/50 | 4.3.2.10 | 312 | 1435 |  | 2520 | 2920 | 3150 | 2770 | 3810 |  |  | NA |  |  |  | 151377530 |  | 1.2 | Milan (3D Printed Plastic)
| 115/50 | 4.3.2.10* | 312 | 1550** |  | 2650 | NA | 3125 | 2750 | 3850 |  |  | NA | [5 empty] | 1.2 | * EU device current ** Extra radio current to investigate
| 120/60 | [11 empty] | NA | [5 empty] | 1.2
[13 empty] | NA | [5 empty] | 1.2
Muse CIC 10 | 110/40 | 4.3.2.10 | 312 | 1250 |  | 2425 | 2850 | 2950 | 2525 | 3625 | NA | NA | NA |  |  |  | 2911322701 / 2911322702 |  | 1.2
| 120/60 | 4.3.2.10* | 312 | 1250 |  | 2450 | NA | 2975 | 2600 | 3725 |  |  | NA |  |  |  | 2911321392 / 2911321391 |  | 1.2 | * EU device current
Muse ITC 312 | 110/40 | [11 empty] | NA | [5 empty] | 1.2
| 115/50 | 4.3.2.10* | 312 | 1350 |  | 2500 | NA | 3250 | 2850 | 3800 | NA | NA | NA |  |  |  | 2911324144 / 2911324145 |  | 1.2 | * EU device current
| 115/50 | 4.3.2.10 | 312 | 1350 |  | 2550 | 2975 | 3025 | 2700 | 3825 | NA | NA | NA |  |  |  | 2911321806 / 2911321807 |  | 1.2
| 120/60 | [11 empty] | NA | [5 empty] | 1.2
[13 empty] | NA | [5 empty] | 1.2
Muse ITE 13 | 110/40 | [11 empty] | NA | [5 empty] | 1.2
| 115/50 | 4.3.2.10 | 13 | 1350 |  | 2475 | 2900 | 3150 | 2725 | 3700 | NA | NA | NA |  |  |  | 2911321364 / 2911321363 |  | 1.2
| 120/60 | 4.3.2.10* | 13 | 1350 |  | 2525 | NA | 3125 | 2700 | 3825 |  |  | NA |  |  |  | 2911322432 / 2911322431 |  | 1.2 | * EU device current
[13 empty] | NA | [5 empty] | 1.2
Gen II 900MHz RIC 312 | 115/50 | 4.9.0.8 | 312 | 1500 |  | 2560 | 2960 | 3090 | 2790 | 3910 | NA | NA | NA |  |  |  | 15613649 |  | 1.2 | *Gen II features ON
| 115/50 | 4.9.0.8 | 312 | 1490 |  | 2650 | 3030 | 350 | 2820 | 3850 | NA | NA | NA |  |  |  | 15613649 |  | 1.2 | *Gen II features OFF
| 115/50 | 4.9.0.8 | 312 | 1530 |  | 2670 | 3120 | 3160 | 2810 | 3890 | NA | NA | NA |  |  |  | 15295637 |  | 1.2 | *Gen II features ON
| 115/50 | 4.9.0.8 | 312 | 1500 |  | 2550 | 2970 | 3200 | 2820 | 3890 | NA | NA | NA |  |  |  | 15295637 |  | 1.2 | *Gen II features OFF
Gen II Sapphire BTE 13 | 130/70 | 4.9.0.8 | 13 | 1410 |  | 2490 | 2910 | 3060 | 2720 | 3810 | NA | NA | NA |  |  |  | 160799038 |  | 1.2 | *Gen II features ON
| 130/70 | 4.9.0.8 | 13 | 1390 |  | 2490 | 2900 | 3050 | 2700 | 3780 | NA | NA | NA |  |  |  | 160799038 |  | 1.2 | *Gen II features OFF
| 130/70 | 4.9.0.8 | 13 | 1400 |  | 2490 | 290 | 3150 | 2770 | 3780 | NA | NA | NA |  |  |  | 160799041 |  | 1.2 | *Gen II features ON
| 130/70 | 4.9.0.8 | 13 | 1390 |  | 2490 | 2900 | 3060 | 2670 | 3790 | NA | NA | NA |  |  |  | 160799041 |  | 1.2 | *Gen II features OFF
Gen II CIC | 120/60 | 4.9.0.8 | 10 | 1.28/1.47* | [12 empty] | 2911321722 |  | 1.2
|  |  |  | SHAoff/SHAon* | [14 empty] | 1.2
Milan Sapphire RIC 312 | 115/50 | 4.10.0.3 | 312 | 1490 |  | 2640 | 3070 | 3120 | 2740 | 1760* | <---- | [5 empty] | 161074534 |  | 1.2 | *US wireless settings
|  | 4.14.0.1 | 312 | 1480 |  | 2620 | 3000 | 3160 | 2770 | 1780* | <---- * In session current | [5 empty] | 161074534 |  | 1.2 | *US wireless settings
|  | 4.14.0.1 | 312 | 1480 |  | NA | 3040 | 3540 | 3110 | 2050* | <---- | [5 empty] | 161074534 |  | 1.2 | *Korean Wireless settings
|  | 4.14.0.3 | 312 | 1450 |  | 2580 | 2980 | 3090 | 2650 | 3780 | 2870 | 2910 |  |  |  |  | 161034891 |  | 1.2
|  | 4.14.0.3 | 312 | 1480 |  | 2590 | 3000 | 3110 | 2780 | 3820 | 2870 | 2850 |  |  |  |  | 161034519 |  | 1.2
Milan Xino RIC 312 | 120/60 | 4.14.0.4 | 312 | 1570 |  | 2810 | 3230 | 3260 | 2850 | 3970 | [6 empty] | 161074534 |  | 1.2 | Z Power rec drain analysis
|  | 4.14.0.4 | 312 | 1610 |  | 2720 | 3150 | 3330 | 2820 | 3980 | [6 empty] | 170007412 |  | 1.2
| 130/70 | 4.14.0.4 | 312 | 1510 |  | 2640 | 3050 | 3350 | 2800 | 3970 | [6 empty] | 170007412 |  | 1.2
Napa BTE13 | 130/70 | 4.7.0.11 | 13 | 1390 | - | 2540 | 2900 | 2970 | 2600 | 3710 | [6 empty] | 161085194 |  | 1.2 | Gen II features on
|  | 4.7.0.11 | 13 | 1400 | - | 2520 | 2920 | 2980 | 2610 | 3710 | [6 empty] | 161085193 |  | 1.2 | Gen II features on
|  | 4.7.0.13 | 13 | 1530 | - | - | - | - | - | 3860 | - | - |  |  |  |  | 161085193 |  | 1.2 | FT on
|  | 4.7.0.18 | 13 | 1379 |  | 2447 | 2851 | 3041 | 2634 | 3677 | [6 empty] | 161085194 |  | 1.2
Napa CIC | 115/50 | 4.7.0.11 | 312 | 1290 | 1380 | 2405 | 2815 | 2820 | 2550 | 3495 | [6 empty] | 2911321720 |  | 1.2
|  | 4.7.0.11 | 312 | 1300 | 1380 | 2451 | 2860 | 2820 | 2540 | 3560 | [6 empty] | 2911321719 |  | 1.2
|  | 4.7.0.13 | 312 | 1430 | 1380 | - | - | - | - | 3650 | - | - |  |  |  |  | 2911321719 |  | 1.2 | FT on
|  | 4.7.0.18 | 10 | 1292 | 1430/1380 | 2430 | 2834 | 2930 | 2540 | 3634 | [6 empty] | 2911323832 |  | 1.2
Napa ITC | 115/50 | 4.7.0.18 | 312 | 1306 | 1435 | 2467 | 2866 | 2987 | 2697 | 3703 | 2864 | 2765
Napa ITE | 130/70 | 4.7.0.18 | 13 | 1568 | 1572 | 2705 | 3105 | 3108 | 2654 | 3737 | 2859 | 2845
Napa Sapphire BTE 312 | 130/60 | 4.7.0.18 | 312 | 1336 | 1449 | 2616 | 3014 | 2969 | 2581 | 3501 | 2853 | 2761
Napa Xino RIC 312 | 115/50 | 4.7.0.18 | 312 | 1439 | 1571 | 2559 | 2966 | 3026 | 2679 | 3731 | 2821 | 2866
Napa Sapphire RIC 312 | 115/50 | 4.7.0.18 | 312 | 1387 | 1548 | 2520 | 2985 | 3112 | 2715 | 3737 | 2850 | 2876 |  |  |  |  | 161234939
161270052
Napa IIC Custom 900MHz | 110/40 | 4.7.0.18 | 10A | 1217 | 1349 | 2358 | 2765 | 2815 | 2553 | 3445 | n/a | n/a
Z series RIC312 AP60 | 119/60AP | 2.15.0.20 | 312 | 1667 | N/A | 3778 | N/A | N/A | N/A | 4279 | N/A | N/A |  |  |  |  | 14044232
Sapphire RIC312 900Mhz | 119/60AP | 4.7.0.18 | 312 | 1641 | 1774 | 2731 | 3130 | 3273 | 2871 | 3964 | [6 empty] | 161235774
|  | 4.7.0.18 | 312 | 1619 | 1751 | 2690 | 3076 | 3233 | 2816 | 3937 | [6 empty] | 170204919
3 Series RIC 312 | 119/60AP | 2.10.4.3 | 312 | 1678 | 1888 | 3698 | N/A | N/A | N/A | N/A | N/A | N/A
Oslo RC RIC 312+ | 115/50 | 4.17.0.3 | 312+ | 559 | 604 | 942 | 1093 | 1140 | 1019 | 1380 | [6 empty] | 170381318 | 3.63 | 1.202 | 3.6 V Device Current Drain
|  |  |  | 1500 | 1660 | 2690 | 3130 | 3265 | 2890 | 4010 | [7 empty] | NA | 1.202 | 1.2 V Current Drain
|  |  |  | 62.3058 | 54.3251 | 51.2617 | 56.5647 | 58.8623 | 62.0358 | 52.1708 | [7 empty] | NA | NA | Calculated PMIC Current Drain
|  | 4.17.0.5 | 312+ | [13 empty] | 170381319 |  |  | 3.6 V Device Current Drain
[18 empty] | NA |  | 1.2 V Current Drain
|  |  |  | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | #DIV/0! | [7 empty] | NA | NA | Calculated PMIC Current Drain
Oslo Sapphire Super Power BTE 13 | 138/80 | 4.7.0.16 | 13 | 1570 |  | 2763 | 3018 | 3236 | 2803 | 3927 | 2915 | 2967 |  |  |  |  | 170381412 |  |  | 90dB brodband stim omni 6654µA Dir 5697µA
|  | 4.7.0.18 | 13 | 1616 | 1753 | 2636 | 3034 | 3181 | 2845 | 3636 | 2812 | 2983 |  |  |  |  | 170714587
|  | 4.18.0.2 | 13 | 1615 | 1739 | 2743 | 3155 | 3176 | 2855 | 3767 | 2833 | 3228
|  | 4.18.1.0 | 13 | 1617 | 1749 | 2712 | 3049 | 3175 | 2816 | 3776 | 2829 | 3023 |  |  |  |  | 170714587
Oslo Xino RIC 312 with 119/60 AP receiver | 119/60AP | 4.7.0.18 | 312 | 1536 | 1670 | 2650 | 3061 | 3141 | 2778 | 3828 | 2798 | 2965
Oslo RC RIC 312+ | 115/50 | 4.17.1.0 | 312+ | 560 | 604 | 924 | 1064 | 1087 | 958 | 1324 | 967 | 945 |  |  |  |  | 171302854 | 3..621 | 1.2001 | with estat BF 50 drain and audio present
|  |  |  | 1507 | 1685 | 2669 | 3072 | 3149 | 2732 | 3757 | 2721 | 2643 | [5 empty] | NA | 1.2
|  | 4.17.1.0 | 312+ | 553 | 605 | 921 | 1054 | 1082 | 960 | 1315 | 956 | 940 |  |  |  |  | 171302853 | 3.623 | 1.2001 | with estat BF 50 drain and audio present
|  |  |  | 1510 | 1676 | 2674 | 3073 | 3112 | 2751 | 3750 | 2715 | 2614 | [5 empty] | NA | 1.2
Oslo RC RIC 312+ | 115/50 | 4.17.1.0 | 312+ | 556 | 608 | 928 | 1058 | 1121 | 972 | 1352 | 1074 | 1043 |  |  |  |  | 171302853 | 3.623
|  |  |  | 1543 | 1662 | 2660 | 3086 | 3171 | 2767 | 3891 | 2919 | 2817 | [6 empty] | 1.2
|  | 4.17.1.0 | 312+ | 536 | 598 | 924 | 1065 | 1114 | 968 | 1367 | 1108 | 1031 |  |  |  |  | 171302854 | 3.621
|  |  |  | 1481 | 1619 | 2632 | 3075 | 3187 | 2765 | 3790 | 3087 | 2903 | [6 empty] | 1.2
|  | 4.17.1.1 | 312+ | 561 | 605 | 908 | 1048 | 1123 | 973 | 1352 | 1035 | 1025 |  |  |  |  | 171302854 | 3.63 | 1.204 | with estat BF 50 drain and audio present
|  |  |  | 1492 | 1623 | 2633 | 3026 | 3209 | 2756 | 3893 | 2987 | 2931
MNR Customs Paris ITE | 130/70 | 4.20.0.2 | 13 | 1455 | 1566 | 2588 | 2992 | 3091 | 2676 | 3832 | DNT | DNT |  |  |  |  | 2911321707
MNR Customs Paris CIC | 110/40 | 4.20.0.2 | 10A | 1312 | 1414 | 2401 | 2805 | 3018 | 2478 | 3516 | DNT | DNT |  |  |  |  | 2911321724

## Charger
Product | Firmware Rev. | Battery Capacity | Supply Voltage (VBAT) | Standby Current (mA) | Charge current, single aid (mA) | Turbo current, single aid (mA) | Charge Current, both aids (mA) | Turbo current, both aids (mA) | Post Charge Current, 1 aid finished (mA) | Post Charge Current 2 aid finished (mA) | Shelf Mode | Expected Charges Size 13 cell | Actual Days Typ Accessory | Expected Days Non Accessory | Expected Charges Size 312+ cell | Serial Numbers | Database ID Code | Notes | NOTE: * = Calculated
Oslo HA Charger | 0.3.072 | 850mAH | 4.2 | 5.89 | 16.7/14.3 | ~ | 27.8/22.8 | ~ | X | 6.4/6.1 | X |  |  |  |  | SN002 (rev1) |  | BK measurement
|  |  | 3.29 | 5.92 | 18.6/17.0 | ~ | 31.0/28.3 | ~ | X | 6.6/6.2 | X | [6 empty] | BK 2
|  |  | 4.2 | 5.92 | 15.4/15.1 | ~ | 25.1/24.6 | ~ | X | X | X |  |  |  |  | SN002 (rev1) |  | Keithley measurement
|  |  | 3.29 | X | 17.6/17.4 | ~ | 29.5/29.2 | ~ | X | X | X | [6 empty] | Keithley measurement
| 0.5.16383 | 850mAH | 4.2 | .860-.540 | 14.8 | ~ | 25.5 | ~ | 15.5 | 5 | X |  |  |  |  | SN002 (rev1)
|  |  | 3.29 | X | X | ~ | X | ~ | X | X | X
| 0.6.0.0 | 850 mAH | 4.2 | 0.531 | 17.7 | ~ | 31.9 | ~ | 18.6 | 5.13 | X |  |  |  |  | rev 2 |  | Size 13 cells
|  |  | 3.6 | 0.668 | 21.7 | ~ | 39.6 | ~ | 24.5 | 9.22 | X
| 0.7.0.0 | 850 mAH | 4.2 | 0.58 | 14.3 | ~ | 23.8 | ~ | 15.3 | 5.86 | 0.0979 |  |  |  |  | rev 2 |  | Size 312 Cells
|  |  | 4 | 0.58 | 15 | ~ | 25.4 | ~ | 16.3 | 6.61 | 0.0935
|  |  | 3.8 | 0.6 | 16 | ~ | 27.8 | ~ | 18 | 7.87 | 0.0894
|  |  | 3.5 | 0.79 | 18.3 | ~ | 32.7 | ~ | 21.9 | 11.1 | 0.0832
| 0.8.0.0 | 850 mAH | 4.2 | 0.56 | 17.9 | ~ | 32.2 | ~ | 18.9 | 5.5 | 0.096 |  |  |  |  | Prod board, Alpha Candidate FW |  | Size 13 Cells
|  |  | 4 | 0.6 | 18.9 | ~ | 34.2 | ~ | 20.4 | 6.39 | 0.092
|  |  | 3.8 | 0.66 | 20.3 | ~ | 36.9 | ~ | 22.5 | 7.85 | 0.088
|  |  | 3.5 | 0.83 | 23.3 | ~ | 42.6 | ~ | 27.2 | 11.29 | 0.082
| 0.8.0.0 | 850 mAH | 4.2 | 0.56 | 14.4 | ~ | 25.1 | ~ | 15.6 | 6.07 | 0.096 |  |  |  |  | Prod board, Alpha Candidate FW |  | Size 312 Cells
|  |  | 4 | 0.6 | 15.2 | ~ | 26.7 | ~ | 16.9 | 7.05 | 0.092
|  |  | 3.8 | 0.66 | 16.5 | ~ | 29 | ~ | 18.8 | 8.64 | 0.088
|  |  | 3.5 | 0.83 | 19.2 | ~ | 33.9 | ~ | 22.9 | 12.25 | 0.082
| 0.10.0.0 | 820 mAh | 4.2 | 0.87 | ~ | ~ | ~ | ~ | ~ | ~ | 0.105 |  |  |  |  | Modified Iterm and Comm R's |  | Size 13 Cells
|  |  | 4 | 0.95 | ~ | ~ | ~ | ~ | ~ | ~ | 0.1
|  |  | 3.8 | 1.05 | ~ | ~ | ~ | ~ | ~ | ~ | 0.096
|  |  | 3.5 | 1.35 | ~ | ~ | ~ | ~ | ~ | ~ | 0.089
| 0.10.0.0 | 820 mAh | 4.2 | 0.87 | 14.3 | ~ | ~ | ~ | ~ | ~ | 0.105 |  |  |  |  | Modified Iterm and Comm R's |  | Size 312 Cells
|  |  | 4 | 0.95 | 15.1 | ~ | ~ | ~ | ~ | ~ | 0.1
|  |  | 3.8 | 1.05 | 16.3 | ~ | ~ | ~ | ~ | ~ | 0.096
|  |  | 3.5 | 1.35 | 18.9 | ~ | ~ | ~ | ~ | ~ | 0.089
| 0.11.0.0 | 820mAh | 4.2 | 0.84 | 14.3 | ~ | 25.2 | ~ | 15.3 | 5.3 | 0.104 | [6 empty] | Size 312 Cells
|  |  | 4 | 0.95 | 15.3 | ~ | 26.5 | ~ | 16.8 | 6.3 | 0.1
|  |  | 3.8 | 1.04 | 16.5 | ~ | 29.7 | ~ | 18.6 | 7.8 | 0.095
|  |  | 3.5 | 1.32 | 19.1 | ~ | 33.5 | ~ | 22.8 | 11.3 | 0.088
| 0.11.0.0 | 820mAh | 4.2 | 0.84 | 18.4 | ~ | 33.2 | ~ | 12.5 | 5.3 | 0.104 | [6 empty] | Size 13 Cells
|  |  | 4 | 0.95 | 19.4 | ~ | 35.1 | ~ | 11.3 | 6.3 | 0.1
|  |  | 3.8 | 1.04 | 20.7 | ~ | 37.6 | ~ | 12.1 | 7.8 | 0.095
|  |  | 3.5 | 1.32 | 23.7 | ~ | 43.2 | ~ | 15 | 11.3 | 0.088
Puck Charger | 2.2.0.0 | 600 mAh | 4.1 | 1.43 | 16.6 | 41 | 29.2 | 75.9 | 7 | 9.6 | 0.225 | NA | X |  | X |  |  | 220 uA Shelf Drain, OSLO CHARGE CURRENTS
|  |  | 3.9 | 1.47 | 17.3 | 42.8 | 30.6 | 79.7 | 7.2 | 9.9 | 0.223
|  |  | 3.7 | 1.49 | 18 | 45 | 32.1 | 83.7 | 7.4 | 10.2 | 0.223
|  |  | 3.5 | 1.5 | 18.9 | 47.4 | 33.8 | 88.7 | 7.7 | 10.5 | 0.223
| 2.2.0.0 | 600 mAh | 4.1 | 1.32 | 16.5 | 40.3 | 28.3 | 75.1 | 7.3 | 10 | 0.125 | NA | X |  | X |  |  | 120 uA Shelf Drain, OSLO CHARGE CURRENTS
|  |  | 3.9 | 1.36 | 17.1 | 42.3 | 29.6 | 78.9 | 7.5 | 10.4 | 0.125
|  |  | 3.7 | 1.4 | 17.8 | 44.4 | 31.1 | 83.1 | 7.7 | 10.8 | 0.125
|  |  | 3.5 | 1.4 | 18.7 | 46.9 | 32.7 | 87.4 | 7.9 | 11.2 | 0.123
| 2.3.0.0 | 600 mAh | 4.1 | 0.47 | 19.5 | 41.7 | 28.4 | 77.3 | 7.1 | 10 | 0.115 | NA | X |  | X |  |  | OSLO CHARGE CURRENTS ONLY
|  |  | 3.9 | 0.46 | 20.3 | 43.8 | 29.8 | 81.4 | 7.2 | 10.3 | 0.115
|  |  | 3.7 | 0.49 | 21.2 | 46.1 | 31.3 | 85.9 | 7.4 | 10.7 | 0.115
|  |  | 3.5 | 0.47 | 22.3 | 48.5 | 33.1 | 90.9 | 7.7 | 11.2 | 0.115
| 2.4.0.0 | 600 mAh | 4.1 | 0.49 | 16.2 | 41.8 | 29.2 | 77.4 | 6.7 | 9.3 | 0.115 | NA |  |  | X |  |  | OSLO CHARGE CURRENTS ONLY
|  |  | 3.9 | 0.49 | 16.9 | 43.8 | 30.7 | 81.4 | 6.9 | 9.6 | 0.115
|  |  | 3.7 | 0.49 | 17.7 | 46 | 32.2 | 85.9 | 7.1 | 10 | 0.115
|  |  | 3.5 | 0.49 | 18.5 | 48.4 | 33.9 | 90.9 | 7.3 | 10.4 | 0.115
| 2.4.0.0 | 600 mAh | 4.1 | ~ | 16.8 | 56.4 | 29.6 | 105.5 | ~ | ~ | ~ | NA |  |  | X |  |  | QUEBEC CHARGE CURRENTS
|  |  | 3.9 | ~ | 17.5 | 59.2 | 31.1 | 111.5 | ~ | ~ | ~
|  |  | 3.7 | ~ | 18.2 | 62.4 | 32.6 | 117.89 | ~ | ~ | ~
|  |  | 3.5 | ~ | 19.1 | 65.9 | 34.4 | 124.9 | ~ | ~ | ~
| 2.5.0.0 | 600 mAh | 4.1 | 0.48 | NC | NC | NC | NC | NC | NC | 0.108 | NA |  |  | No Change |  |  | OSLO CHARGE CURRENTS ONLY
|  |  | 3.9 | 0.48 | [6 empty] | 0.108
|  |  | 3.7 | 0.48 | [6 empty] | 0.108
|  |  | 3.5 | 0.48 | [6 empty] | 0.108
| 2.5.0.0 | 600 mAh | 4.1 | 0.48 | 16.9 | 56.5 | 29.4 | 106 | 6.6 | 9.2 | 0.108 | NA |  |  | No Change |  |  | QUEBEC CHARGE CURRENTS
|  |  | 3.9 | 0.48 | 17.6 | 59.4 | 30.8 | 11.7 | 6.8 | 9.5 | 0.108
|  |  | 3.7 | 0.48 | 18.4 | 62.5 | 32.4 | 118 | 6.9 | 9.9 | 0.108
|  |  | 3.5 | 0.48 | 19.3 | 66.1 | 34.2 | 125.2 | 7.1 | 10.3 | 0.108
| 2.7.0.0 | 600 mAh | 4.1 | 0.48 | NC | NC | NC | NC | NC | NC | NC | NA |  |  | No Change
|  |  | 3.9 | 0.49
|  |  | 3.7 | 0.48
|  |  | 3.5 | 0.48

## 2.4 GHz + NFMI
Product | Matrix | Firmware Rev. | Battery Size | Advertising Current (µA) | 30ms Sync Current (µA) | iPhone Stream Current (µA) | iPhone Call Current (µA) | SoundSpace Current (µA) | Stream+ SoundSpace Current (µA) | Nominal Current (µA) | E2E Transmit Current (µA) | E2E Receive Current (µA) | TVS Stream Current (µA) | MFA Stream Current (µA) | MFA Phone Call Current (µA) | BNR Stream Current (µA) | MNR + Stimulus Current (µA) | 90 dBSPL Stimulus Current (µA) | Slow Advertisement Current (µA) | Android connection Current (uA) | Android Streaming Current (uA) | Expected Days Typ Accessory | Actual Days Typ Accessory | Expected Days Non Accessory | Actual Days Non Accessory | Serial Numbers | Database ID Code | Notes | Date | NOTE: * = Calculated
Parlo BTE13 | 130/70 | 6.0.0.7 | 13 | 3590 | NA |  |  |  |  | 2320 | [15 empty] | 171331841
Paris BTE13 PIG flagship | 130/70 | 6.0.1.3 | 13 | 3360 | [5 empty] | 2568 | [15 empty] | 171349304 |  | This aid seemed ok, then 06 device's nominal stayed up and now both aids read that way now.
|  | 6.0.1.3 | 13 | 3298 | [5 empty] | 3125 | [15 empty] | 171349306
Paris RIC312 Pig flagship | 115/50 | 6.0.1.4 | 312 | 3323 | [5 empty] | 3104 | 4270 E2E stream phone | [14 empty] | 180005573
| 115/51 | 6.0.1.4 | 312 | 3235 | [5 empty] | 3140 | 3791 E2E Stream receive | [14 empty] | 180005570
Paris BTE13 PIG flagship | 130/70 | 6.0.1.5 | 13 | 5136 | [5 empty] | 4673 | [15 empty] | 171146846 |  | Radio is always on.
|  | 6.0.1.5 | 13 | 5268 | [5 empty] | 4715 | [15 empty] | 171146849 |  | Radio is always on.
Paris RIC312 Pig flagship | 115/50 | 6.0.1.5 | 312 | 5612 | [5 empty] | 5253 | [15 empty] | 180005570
| 115/50 | 6.0.1.5 | 312 | 5543 | [5 empty] | 5124 | [15 empty] | 180005573
Paris BTE13 Pig Flagship | 130/70 | 6.0.1.8 | 13 | 2680 | 2610 | 3333 | 3084 |  |  | 2558 | 4157 | 3782 |  |  |  | 4228 | 2494 |  | 2420 | [6 empty] | 171349304
|  | 6.0.1.8 | 13 | 2780 | 2710 | 3414 | 3202 |  |  | 2667 | 4225 | 3807 |  |  |  | 4420 | 2666 |  | 2591 | [6 empty] | 171146852
Paris RIC312 Pig Flagship | 115/50 | 6.0.1.8 | 312 | 2897 | 2820 | 3488 | 3451 |  |  | 2756 | 4438 | 4104 |  |  |  | 4342 | 2553 |  | 2553 | [6 empty] | 171302871
|  | 6.0.1.8 | 312 | 2965 | 2916 | 3549 | 3541 |  |  | 2829 | 4244 | 3812 |  |  |  | 4414 | 2642 |  | 2642 | [6 empty] | 171302873
Paris RIC312 Pig Flagship | 115/50 | 6.0.1.9 | 312 | 2170 | 2170 | 3350 | 3360 |  |  | 2170 | 3570 | 3330 |  |  |  | 3520 | 1940 |  | 1930 | [6 empty] | 171302873 |  | Standby Mode was enabled.
|  | 6.0.1.9 | 312 | 2150 | 2060 | 3290 | 3270 |  |  | 2090 | 3520 | 3240 |  |  |  | 3680 | 2080 |  | 1834 | [6 empty] | 171302871 |  | Standby Mode was enabled.
Paris BTE13 Pig Flagship | 130/70 | 6.0.1.9 | 13 | 2120 | 2260 | 3127 | DNT |  |  | 2050 | 3180 | 3000 |  |  |  | 3570 | 1940 |  | 1930 | [6 empty] | 180367245 |  | Because of wireless connectivity issues, the BNR, MNR, E2E Phone current drains were measured without a phone connected. iPhone Call current wasn't measured due to drop outs and buzzing. Standby Mode was enabled.
|  | 6.0.1.9 | 13 | 2130 | 2300 | 3133 | DNT |  |  | 2060 | 3240 | 3060 |  |  |  | 3560 | 1960 |  | 1940 | [6 empty] | 180367242
Paris RIC312 Panda Flagship | 115/50 | 6.0.1.11 | 312 | 2167 | 2279 | 3162 | 3157 |  |  | 2049 | 3392 | 3222 |  |  |  | 3750 | 1975 |  | 1975 | [6 empty] | 180367120
|  | 6.0.1.11 | 312 | 2197 | 2261 | 3155 | 3132 |  |  | 2179 | 3473 | 3346 |  |  |  | 3740 | 1925 |  | 1925 | [6 empty] | 180367336
Paris BTE13 Panda Flagship | 130/70 | 6.0.1.11 | 13 | 2099 | 2286 | 3182 | 3142 |  |  | 2116 | 3472 | 3185 |  |  |  | 3710 | 2195 |  | 1922 | [6 empty] | 180367242
|  | 6.0.1.11 | 13 | 2088 | 2245 | 3147 | 3127 |  |  | 2015 | 3492 | 3284 |  |  |  | 3815 | 2153 |  | 1926 | [6 empty] | 180367245
Paris BTE13 Panda Flagship | 130/70 | 6.0.1.12 | 13 | 2230 | 2110 | 3130 | 3150 |  |  | 1980 | 3410 | 3070 |  | 2820 |  | 3550 | 1980 |  | 1860 | [6 empty] | 180367242
|  | 6.0.1.12 | 13 | 2150 | 2090 | 3110 | 3130 |  |  | 1960 | 3360 | 3170 |  | 2790 |  | 3610 | 1960 |  | 1870 | [6 empty] | 180367245
Paris BTE13 Panda Flagship | 130/70 | 6.0.1.13 | 13 | 2210 | 2260 | 3160 | 3140 |  |  | 1940 | 3380 | 3070 | 2820 |  |  | 3490 | 2020 |  | 1830 | [6 empty] | 180367242
|  | 6.0.1.13 | 13 | 2140 | 2270 | 3210 | 3140 |  |  | 1980 | 3520 | 3320 | 2850 |  |  | 3690 | 2090 |  | 1820 | [6 empty] | 180367245
Paris RIC312 Panda Flagship | 115/50 | 6.0.1.13 | 312 | 2150 | 2110 | 3240 | 3280 |  |  | 2030 | 3450 | 3130 | 2940 |  |  | 3630 | 2070 |  | 1970 | [6 empty] | 180387149
|  | 6.0.1.13 | 312 | 2120 | 2110 | 3220 | 3230 |  |  | 2020 | 3410 | 3190 | 2980 |  |  | 3650 | 2060 |  | 1910 | [6 empty] | 180387153
Paris BTE13 Panda Flagship | 130/70 | 6.0.1.15 | 13 | 2110 | 2120 | 3190 | 3180 |  |  | 2120 | 3460 | 3230 | 2880 |  |  | 3790 | 2120 |  | 1930 | [6 empty] | 180367245
|  | 6.0.1.15 | 13 | 2120 | 2140 | 3190 | 3180 |  |  | 2100 | 3480 | 3290 | 2870 |  |  | 3690 | 2110 |  | 1940 | [6 empty] | 180367242
Paris BTE13 Panda Flagship | 130/70 | 6.0.2.1 | 13 | 2023 | 2050 | 3049 | 3095 |  |  | 1993 | 3794 | 3216 | 2718 | 2697 |  | 3667 | 2124 |  | 2051 | [6 empty] | 180913673
|  | 6.0.2.1 | 13 | 2064 | 2093 | 3073 | 3105 |  |  | 1839 | 3689 | 3209 | 2752 | 2758 |  | 3413 | 2194 |  | 1957 | [6 empty] | 180581681
Paris BTE13 Panda Flagship | 130/70 | 6.0.2.2 | 13 | 2017 | 2182 | 3124 | 3150 |  |  | 1792 | 3108 | 2956 | 2734 | 2724 |  | 3456 | 2187 |  | 1816
|  | 6.0.2.2 | 13 | 2057 | 2139 | 3119 | 3042 |  |  | 1839 | 3048 | 2785 | 2738 | 2754 |  | 3425 | 1975 |  | 1945
Paris RIC312 Panda Flagship | 115/50 | 6.0.2.2 | 312 | 2150 | 2120 | 3130 | 3130 |  |  | 2050 | 3370 | 3120 |  | 2850 |  | 3550 | 2030 |  | 1890 | [6 empty] | 180367110
|  | 6.0.2.2 | 312 | 2170 | 2120 | 3140 | 3130 |  |  | 2070 | 3370 | 3120 |  | 2860 |  | 3620 | 2040 |  | 1890 | [6 empty] | 180367111
Paris RIC312 Panda Flagship | 115/50 | 6.0.2.3 | 312 | 2160 | 2130 | 3105 |  |  |  | 2010 | [5 empty] | 3550 |  |  | 1920 | [6 empty] | 180367110 |  | Right - iPhone SE
|  | 6.0.2.3 | 312 | 2240 | 2170 | 3170 |  |  |  | 2060 | [5 empty] | 3600 |  |  | 1970 | [6 empty] | 180367111 |  | Left - iPhone SE
Quebec RIC312+ no double tap | 115/50 | 6.3.0.5 | 312+ | 872 | 907 | 1197 |  |  |  | 818 | [8 empty] | 807 | [6 empty] | 170381318 | 3.8 | 1.253 | 3.8 V Device Current Drain
|  |  |  | 2560 | 2580 | 3670 |  |  |  | 2470 | 3950 | 3730 | 3390 |  |  | 4150 | 2373 |  | 2350 | [7 empty] | NA | 1.253 | 1.25 V Current Drain
|  |  |  | 27.8737 | 56.2789 | -13.1342 | 0 | 0 | 0 | 3.55 | -1302.4605 | -1229.9184 | -1117.8079 | 0 | 0 | -1368.4079 | -782.4655 |  | 32.1184 | [7 empty] | NA | NA | Calculated PMIC Current Drain
Quebec WiCROS Quokka RIC REC | 115/50 | 6.3.0.8 | 312 | 3410 | 3520 | 3760 | Native app not pairing to the pair |  |  | 3170 | N/A | N/A | 3500 |  |  | N/A | 3610 | 3550 | 3150 | [6 empty] | 180367153
Quebec WiCROS Quokka RIC Trans | 115/50 | 6.3.0.8 | 312 | 3140 | N/A | N/A | Native app not pairing to the pair |  |  | 2880 | N/A | N/A | 3100 |  |  | N/A | 3310 | 3300 | 2750 | [6 empty] | 180367149
Quebec RIC312+ Queen | 115/50 | 6.3.0.8 | 312 | 2130 |  | 3090 |  |  |  | 2000 | 3650 | 3270 |  | 2810 |  | 3830 |  |  | 1910 | [6 empty] | 181054783
Quebec WiCROS Queen RIC REC | 115/50 | 6.3.0.8 | rechargeable | 3730 | 3680 | 3810 | 3730 |  |  | 1760 | N/A | N/A |  | 3350 |  |  |  |  | 3470 | [6 empty] | 181054783
Quebec WiCROS Queen RIC TRANS | 115/50 | 6.3.0.8 | rechargeable | 3310 | N/A | N/A |  |  |  | 2390 | N/A | N/A | [6 empty] | 3060 | [6 empty] | 181054784
Quebec Queen RRIC | 115/50 | 6.3.0.8 | rechargeable | 2110 | 2150 | 3120 | 3070 | Thrive crashing | Thrive crashing | 1910 | 3630 | 3510 | 2900 | 2830 |  | 3860 | 2170 | [8 empty] | 181054783
Quebec Queen RRIC | 115/50 | 6.3.0.8 | rechargeable | 2130 | 2170 | 3140 | 3110 | Thrive crashing | Thrive crashing | 1930 | 3590 | 3480 | 2860 | 2820 |  | 3890 | 2180 | [8 empty] | 181054784
Quebec Queen RRIC | 115/50 | 6.3.1.4 | rechargeable | 2220 | 2180 | 3160 | 3150 |  |  | 2150 | 3810 | 3350 |  | 2900 |  | 4040 | 2140 |  | 1960 | [6 empty] | 181043497 |  | iPhone XR and MFA
Quebec Queen RRIC | 115/50 | 6.3.1.4 | rechargeable | 2240 | 2180 | 3150 | 3160 |  |  | 2160 | 3900 | 3520 |  | 2900 |  | 4110 | 2180 |  | 1990 | [6 empty] | 181043496 |  | iPhone XR and MFA
Quebec WiCROS Queen RIC REC | 115/50 | 6.3.1.4 | rechargeable | 3850 | 3780 | 3630 | 3640 |  |  | 3790 | N/A | N/A |  | 3350 |  | 3780 | 3790 |  | 3580 | [6 empty] | 181043496 |  | iPhone XR and MFA
Quebec WiCROS Queen RIC Trans | 115/50 | 6.3.1.4 | rechargeable | 3550 | 3250 | 2150 | 2140 |  |  | 3270 | N/A | N/A |  | 2150 |  | 3200 | 3240 |  | 3190 | [6 empty] | 181043497 |  | iPhone XR and MFA
Quebec Queen RRIC | 115/50 | 6.3.1.5 | rechargeable | 2300 | 2040 | 3120 | 3140 |  |  | 2030 | 3740 | 3310 |  | 2770 |  | 3870 | 2030 |  | 1960 | [6 empty] | 181043497 |  | iPhone XR and MFA
Quebec Queen RRIC | 115/50 | 6.3.1.5 | rechargeable | 2300 | 2030 | 3130 | 3150 |  |  | 2030 | 3730 | 3330 |  | 2770 |  | 3960 | 2030 |  | 1980 | [6 empty] | 181043496 |  | iPhone XR and MFA
Quebec WiCROS Quoka RIC REC | 115/50 | 6.3.1.4 |  | 3840 | 3450 | 3580 |  |  |  | 3420 | N/A | N/A |  | 3180 |  | N/A |  |  | 3370 | [6 empty] | 180367110 |  | iPhone SE and MFA
Quebec WiCROS Quoka RIC Trans | 115/50 | 6.3.1.4 |  | 3240 | 2940 | 2010 |  |  |  | 2920 | N/A | N/A |  | 1980 |  | N/A |  |  | 2920 | [6 empty] | 180367111 |  | iPhone SE and MFA - Continues to Fast advertise for two minutes
Paris RIC312 Quoka Flagship | 115/50 | 6.3.1.5 | 312 | 2300 | 2050 | 3170 |  |  |  | 2000 | 3380 | 3130 | 2850 |  |  | 3590 | 2040 |  | 1990 | [6 empty] | 180367110 |  | iPhone SE and TV Streamer - Right
|  | 6.3.1.5 | 312 | 2250 | 2060 | 3200 |  |  |  | 2020 | 3410 | 3170 | 2870 |  |  | 3570 | 2020 |  | 2000 | [6 empty] | 180367111 |  | iPhone SE and TV Streamer - Left
Quebec Queen plus Rc RIC | 115/50 | 6.3.2.3 | rechargeable | 2300 | 2160 | 3130 | 3120 | iOS blocked | blocked | 1910 | 3690 | 3530 | 2600 | 2630 |  | 3770 | 2640 | 2250 | 1850 | [6 empty] | 181054596 |  | iPhone 5S (12.0), Thrive build 107, | 1.20 Vbat
Quebec Quokka Plus BTE | 130/70 | 6.3.3.5 | 13 | 2180 | 1980 | 3120 |  |  |  | 1980 | 3430 | 3130 |  |  |  | 3570 | 1970 |  | 1900 | [6 empty] | 180367242 |  | iPhone 6s
|  | 6.3.3.5 | 13 | 2310 | 2060 | 3150 |  |  |  | 1970 | 3350 | 3110 |  |  |  | 3570 | 2010 |  | 1920 | [6 empty] | 180367245 |  | iPhone 6s
Quebec Quokka Plus RIC | 115/50 | 6.3.3.5 | 312 | 2290 | 2070 | 3230 |  |  |  | 1980 | 3560 | 3220 |  |  |  | 3730 | 2060 |  | 1990 | [6 empty] | 180367130 |  | iPhone 6s
|  | 6.3.3.5 | 312 | 2330 | 2070 | 3220 |  |  |  | 2030 | 3460 | 3250 |  |  |  | 3630 | 2020 |  | 1940 | [6 empty] | 180581004 |  | iPhone 6s
Paris RIC + Quebec FW | 115/50 | 6.3.4.1 | 312 | 2180 | 2090 | 3130 |  |  |  | 2030 | 3450 | 3120 |  | 2880 |  | 3700 | 2040 |  | 1960 | [6 empty] | 180581004 |  | iPhone 6s
|  | 6.3.4.1 | 312 | 2200 | 2100 | 3210 |  |  |  | 2060 | 3480 | 3240 |  | 2890 |  | 3750 | 2050 |  | 1940 | [6 empty] | 180367130 |  | iPhone 6s
Paris BTE + Quebec FW | 130/70 | 6.3.4.1 | 13 | 2120 | 1960 | 3110 |  |  |  | 1970 | 3420 | 3170 |  | 2800 |  | 3600 | 1980 |  | 1910 | [6 empty] | 180367245 |  | iPhone 7+
|  | 6.3.4.1 | 13 | 2120 | 1980 | 3130 |  |  |  | 1990 | 3360 | 3040 |  | 2820 |  | 3460 | 1970 |  | 1890 | [6 empty] | 180367242 |  | iPhone 7+
Quebec R RIC | 115/50 | 6.3.4.1 | Rechargeable | 2130 | 2050 | 3140 |  |  |  | 2030 | [6 empty] | 1990 | [8 empty] | 181054601
Rome Quebec RC RIC |  | 6.5.0.8 | Rechargeable | 2164 | 2043 | 3127 | 3134 |  |  | 1953 | 3721 | 2014 | 2782 | 2834 |  |  |  | 2142 | 1953 | [6 empty] | 181054601
R02 - Vbat domain | 115/50 | 6.5.1.6 | 312+ | 680 | 640 | 1030 | 1020 | 870 |  | 640 | 1020 | 1170 |  | 910 |  | 1210 |  | 1100 | [7 empty] | 191323934
R02 - VHA domain (calculated) |  |  |  | 1860 | 1740 | 2910 | 2880 | 2430 |  | 1740 | 2880 | 3330 |  | 2550 |  | 3450 |  | 3120
Rome BTE - Vbat domain | 130/70 | 7.1.2.1 | 312+ | 730 | 660 | 1010 |  |  |  | 650 | [8 empty] | 630 | [6 empty] | 190645351 |  | iPhone SE, Double Tap Enabled, VHA=1.2, VBAT=3.6
Rome BTE - VHA domain (calculated) |  |  |  | 1920 |  | 2850 |  |  |  | 1800 | [8 empty] | 1710
Rome BTE - Vbat domain | 130/70 | 7.1.2.1 | 312+ | [22 empty] | 190645351 |  | 32 kHz Crystal Off
Rome BTE - VHA domain (calculated) |  |  |  | 1940 |  | 2880 |  |  |  | 1820 | [8 empty] | 1740
Rome BTE - Vbat domain | 130/70 | 7.1.2.1 | 312+ | [22 empty] | 190645351 |  | Auto ON/OFF Enabled
Rome BTE - VHA domain (calculated) |  |  |  | 2020 |  | 2930 |  |  |  | 1890 | [8 empty] | 1810
Paris BTE + Tulum FW | 130/70 | 6.10.0.4 | 13 | 2113 | 1919 | 3019 | 3035 |  |  | 1927 | 3306 | 3030 |  | 2795 |  | 3517 |  |  | 1887 | [6 empty] | 200208926 |  | Auto ON/OFF and Double Tap Enabled
|  |  | 13 | 2074 | 1891 | 2963 | 2955 |  |  | 1896 | 3300 | 3011 |  | 2713 |  | 3475 |  |  | 1852 | [6 empty] | 200208921 |  | Auto ON/OFF and Double Tap Enabled
Paris BTE + Tulum FW | 130/70 | 6.10.0.9 | 13 | 1980 | 1790 | 2840 |  |  |  | 1810 | [15 empty] | 190673364 |  | iPhone 7 (13.2)
|  |  | 13 | 2010 | 1810 | 2890 |  |  |  | 1830 | [15 empty] | 190673365 |  | iPhone 7 (13.2)
Quebec RRIC + Tulum FW IP16 VHA | 115/50 | 6.10.0.9 | RC | 2150 | 1980 | 3110 |  |  |  | 1940 | [15 empty] | 20088358 |  | iPhone 7 (13.2)
Model S3 RIC312 | 120/60 | 6.11.0.6 | 312 | 2150 | N/A | N/A | N/A | N/A | N/A | 2070 |  |  | 2780 | 2810 |  |  | 2070 | N/A | N/A | [6 empty] | 201684781 |  | Model S devices are unable to pair w/ phone. Had to skip all the phone measurements | 2021-05-20
| 120/60 |  | 312 | 2160 | N/A | N/A | N/A | N/A | N/A | 2070 |  |  | 2780 | 2810 |  |  | 2070 | N/A | N/A | [6 empty] | 201684786
Model S3 BTE13 | 130/70 | 6.11.0.6 | 13 | 2050 | N/A | N/A | N/A | N/A | N/A | 1750 |  |  | 2610 | 2650 |  |  | 1770 | N/A | N/A | [6 empty] | 190673360 |  | Model S devices are unable to pair w/ phone. Had to skip all the phone measurements | 2021-05-20
| 130/70 |  | 13 | 2050 | N/A | N/A | N/A | N/A | N/A | 1750 |  |  | 2610 | 2650 |  |  | 1770 | N/A | N/A | [6 empty] | 190673361
Utica 2.4GHz Custom CIC | 130/70 | 7.5.0.8 | 312 | 2300 | 1870 | 2920 | 3750 | 2090 | 3150 | 1930 | N/A | N/A | 2790 | 2770 |  | N/A | 2440 | 2210 | 1750 | [6 empty] | 2911330919 |  |  | 2021-06-14
| 130/70 |  | 312 | 2290 | 1910 | 2926 | 2960 | 2080 | 3170 | 1960 | N/A | N/A | 2780 | 2810 |  | N/A | 2490 | 2180 | 1760 | [6 empty] | 2911330918
Utica PPBTE | 130/70 | 7.5.0.8 | 13 | 2365 | 3263 | 3452 | 3445 | 2307 | 3682 | 2167 | N/A | N/A | 3135 | 3138 | 3192 | N/A | 2446 | 4968 | 2071 | [6 empty] | 201535553 |  |  | 2021-07-13
| 130/70 |  | 13 | 2467 | 3235 | 3587 | 3552 | 2451 | 3723 | 2379 | N/A | N/A | 3247 | 3285 | 3279 | N/A | 2428 | 5053 | 2123 | [6 empty] | 201535552
Utica cont. 1 (Evolv AI) RIC 312 | 115/50 | 6.11.1.0 | 312 | 2110 | 2050 | 2710 | 3120 | 1980 | 2790 | 1960 | 3450 | 3240 | 2750 | 2730 | N/A | 2640 | 1940 |  |  |  | 2890 |  |  |  |  | 201684781 |  |  | 2021-09-17
| 115/50 |  | 312 | 2120 | 1950 | 2720 | 3130 | 1970 | 2810 | 1940 | 3480 | 3160 | 2760 | 2740 | N/A | 2840 | 1920 |  |  |  | 2850 |  |  |  |  | 201684786
Xanadu Energizer R02D RRIC | 115/50 | 6.16.0.6 | Battery current | 902 max | 756 | 1016 |  |  |  | 690 | 1250 | 1120 |  | 1018 | [5 empty] | 740 | 1020 |  |  |  |  | 212182325 |  | DT disabled | 2023-01-06
Evolv AI 2400 |  |  | Vha current | 2706 | 2268 | 3048 |  |  |  | 2070 | 3750 | 3360 |  | 3054 | [5 empty] | 2220 | 3060
Note on A3: Seyed Pourbakhsh: Data collected from old 2.4 GHz current drain test procedure.
Note on A5: Seyed Pourbakhsh: Data collected from old 2.4 GHz current drain test procedure.
Note on K5: Bob Rusten: After 06's high Nominal now both aids stay near advertising drain.
Note on A8: Seyed Pourbakhsh: Data collected from old 2.4 GHz current drain test procedure.
Note on A11: Seyed Pourbakhsh: Data collected from old 2.4 GHz current drain test procedure.
Note on A14: Seyed Pourbakhsh: Data collected from old 2.4 GHz current drain test procedure.

## 2.4 GHz
Product | Matrix | Firmware Rev. | Battery Size | Advertising Current (µA) | 30ms Sync Current (µA) | iPhone Stream Current (µA) | iPhone Call Current (µA) | SoundSpace Current (µA) | Stream+ SoundSpace Current (µA) | Nominal Current (µA) | TVS Stream Current (µA) | MFA Stream Current (µA) | MFA Phone Call Current (µA) | 90 dBSPL Stimulus Current (µA) | Stream Current (µA) | Stream+ Soundpoint Current (µA) | Slow Advertisement Current (µA) | Android connection Current (uA) | Android Streaming Current (uA) | Auto Sleep Current (uA) | Expected Days Typ Accessory | Actual Days Typ Accessory | Expected Days Non Accessory | Actual Days Non Accessory | Serial Numbers | Database ID Code | Notes | Date | NOTE: * = Calculated
London RIC 13 | 115/50 | 5.0.1.6.275 | 13 |  |  | 5450 |  |  |  | 2150 |  |  |  |  | 5450 | [5 empty] | 7 | *6.75 | NA | NA | 15797890 / 15797891 |  | Typical Wireless Accessory
|  | 5.0.1.7 | 13 |  |  | 5425 |  |  |  | 2125 |  |  |  |  | 5425 | [5 empty] | 7 | 6.8 | NA | NA | 15797890 / 15797891 | A2151 | Typical Wireless Accessory Original Profile
|  | 5.0.1.7 | 13 |  |  | 5425 |  |  |  | 2125 |  |  |  |  | 5425 | [5 empty] | 7 | 7.25 | NA | NA | 15797890 / 15797891 | A2154 | Typical Wireless Accessory New Profile 14 Hr Day
|  | 5.0.1.8 | 13 |  |  | 5375 |  |  |  | 2100 |  |  |  |  | 5375 | [5 empty] | 7 | 7.25 | NA | NA | 15797892 / 15797893 | A2154 | Typical Wireless Accessory New Profile 14 Hr Day
|  | 5.0.1.9 | 13 |  |  | 5475 |  |  |  | 2000 |  |  |  |  | 5475 | [5 empty] | 7 |  | NA | NA | 15797892 / 15797893
|  | 5.0.1.10 | 13 |  |  | 5450 |  |  |  | 1725 |  |  |  |  | 5450 | [5 empty] | 7 | *8.10 | NA | NA | 15908227 / 15797887
|  | 5.0.1.12 | 13 |  |  | 5525 |  |  |  | 1725 |  |  |  |  | 5525 | [5 empty] | 7 | *8.10 | NA | NA | 15908227 / 15797887
|  | 5.0.2.5 | 13 |  |  | 5500 |  |  |  | 1750 |  |  |  |  | 5500 | [5 empty] | 7 | *8.10 | NA | NA | 15908228 / 1590824
|  | 5.0.2.7 | 13 |  |  | 5500 |  |  |  | 1725 |  |  |  |  | 5500 | [5 empty] | 7 | 7.25 | NA | NA | 15295598 |  | Tested with Gen4 evaluation batteries
|  | 5.0.2.7 | 13 |  | 2430 | 5450 |  |  |  | 1730 |  |  |  |  | 5450 | [9 empty] | 15295599
|  | 5.0.2.7 | 13 |  | 2200 | 5450 |  |  |  | 1720 |  |  |  |  | 5450 | [9 empty] | 1597886
|  | 5.0.2.7 | 13 |  | 2220 | 5480 |  |  |  | 1730 |  |  |  |  | 5480 | [9 empty] | 15295600
|  | 5.0.2.8 | 13 | 2460 | 2160 | 5400 |  | Not available |  | 1650 |  |  |  |  | 5400 | Not available | [8 empty] | 170259151 |  | WK - 6-16-17
|  | 5.0.2.8 | 13 | 2470 | 2210 | 5370 |  | Not available |  | 1720 |  |  |  |  | 5370 | Not available | [8 empty] | 160502929 |  | WK - 6-16-17
|  | 5.6.0.1 | 13 |  | 2180 | 5420 |  |  |  | 1740 |  |  |  |  | 5420 | [9 empty] | 15295599
|  | 5.6.0.6 | 13 | 2470 | 2140 | 5500 |  | 7200 |  | 1650 |  |  |  |  | 5500 | 7200 | [8 empty] | 170259151 |  | WK - 6-15-17
|  | 5.6.0.6 | 13 | 2440 | 2140 | 5430 |  | 7220 |  | 1670 |  |  |  |  | 5430 | 7220 | [8 empty] | 160502929 |  | WK - 6-15-17
| 120/60 | 5.0.1.8 | 13 |  |  | 5400 |  |  |  | 2200 |  |  |  |  | 5400 | [5 empty] | 6 | *6.90 | NA | NA | 15797892 / 15797893
|  | 5.0.1.10 | 13 |  |  | 5600 |  |  |  | 1800 |  |  |  |  | 5600 | [5 empty] | 6 | *7.75 | NA | NA | 15908227 / 15797887
|  | 5.0.1.12 | 13 |  |  | 5600 |  |  |  | 1825 |  |  |  |  | 5600 | [5 empty] | NA |  |  |  | 15908227 / 15797887
| 123/60 | 5.0.1.10 | 13 |  |  | 5550 |  |  |  | 1850 |  |  |  |  | 5550 | [5 empty] | NA |  |  |  | 15908227 / 15797887
|  | 5.0.1.12 | 13 |  |  | 5450 |  |  |  | 1700 |  |  |  |  | 5450 | [5 empty] | NA |  |  |  | 15908227 / 15797888
| 130/70 | 5.0.1.8 | 13 |  |  | 5400 |  |  |  | 2125 |  |  |  |  | 5400 | [5 empty] | 6 | *7.00 | NA | NA | 15797892 / 15797893
|  | 5.0.1.10 | 13 |  |  | 5550 |  |  |  | 1725 |  |  |  |  | 5550 | [5 empty] | 6 | *8.10 | NA | NA | 15908227 / 15797887
|  | 5.0.1.12 | 13 |  |  | 5550 |  |  |  | 1750 |  |  |  |  | 5550 | [5 empty] | 6 |  |  |  | 15908227 / 15797887
GnResound Lynx2 | 119/59 |  |  |  |  | 4900 |  |  |  | 2375 |  |  |  |  | 4900 | [5 empty] | NA | *4.50 | NA | NA
[6 empty] | 4400 | [8 empty] | 4400
Milan RIC 312 | 115/50 | 5.0.2.7 | 312 |  | 2256 | 5692 |  |  |  | 1434 |  |  |  |  | 5692 | [9 empty] | 160127486
|  | 5.6.0.2 | 312 |  | 2350 | 5410 |  |  |  | 1690 |  |  |  |  | 5410 | [9 empty] | 161034569
|  | 5.6.0.2 | 312 |  |  | 5500 |  | 8500 | [6 empty] | 5500 | 8500 | [8 empty] | unknown
|  | 5.6.0.6 | 312 | 2250 | 2110 | 5380 |  | 7300 |  | 1700 |  |  |  |  | 5380 | 7300 | [8 empty] | 161307865 |  | WK - 6-16-17
|  | 5.6.0.6 | 312 | 2350 | 2100 | 5320 |  | 7270 |  | 1620 |  |  |  |  | 5320 | 7270 | [8 empty] | 161307858 |  | WK - 6-19-17
176
Napa Ruby RIC13 | 115/50 | 5.7.0.9 | 13 | 2510 | 2380 (760ms) | 5550 |  |  |  | 1790 |  |  |  |  | 5550 | [9 empty] | 170259151 |  | 460 ms long ad rate would add 80uA to any current except Advertising
|  | 5.7.0.17 | 13 | 2394 | 2404 | 5454 |  |  |  | 1535 |  |  |  |  | 5454 |  |  |  |  | 184 |  |  |  |  | 170255917 |  | 2017-08-01
Napa Sapphire RIC 312 | 115/50 | 5.7.0.10 | 312 |  |  | 5550 |  | 6500 | [6 empty] | 5550 | 6500 | [8 empty] | 170259151
|  | 5.7.0.12 | 312 | 2440 | 2250 (760) | 5550 |  |  |  | 1700 |  |  |  |  | 5550 |  |  |  |  | 185 |  |  |  |  | 161034566
|  | 5.7.0.17 | 312 | 2266 | 2366 | 5546 |  |  |  | 1647 |  |  |  |  | 5546 | [9 empty] | 161307872
Oslo Sapphire RIC 312 | 119/60 | 5.7.0.17 | 312 |  | 2.32mA | 5.59mA |  |  |  | 1.90mA |  |  |  |  | 5.59mA | [9 empty] | 161307851
Oslo Sapphire RIC 312 | 119/60 | 5.7.0.17 | 312 |  | 2.39mA | 5.58mA |  |  |  | 1.94mA |  |  |  |  | 5.58mA | [9 empty] | 161034569
Oslo IP6 RIC13 | AP 119/60 | 5.7.0.17 | 13 | 2480 | 2454 | 5635 |  |  |  | 1945 |  |  |  |  | 5635 | [9 empty] | 15295600
Parlo BTE13 | 130/70 | 6.0.0.7 | 13 | 3590 | NA | NA |  |  |  | 2320 |  |  |  |  | NA | [9 empty] | 171331841
Paris BTE13 PIG flagship | 130/70 | 6.0.1.3 | 13 | 3360 | [5 empty] | 2568 | [14 empty] | 171349304 |  | This aid seemed ok, then 06 device's nominal stayed up and now both aids read that way now.
|  | 6.0.1.3 | 13 | 3298 | [5 empty] | 3125 | [14 empty] | 171349306
Paris RIC312 Pig flagship | 115/50 | 6.0.1.4 | 312 | 3323 |  | 4270 E2E stream phone |  |  |  | 3104 |  |  |  |  | 4270 E2E stream phone | [9 empty] | 180005573
Paris RIC312 Pig flagship | 115/51 | 6.0.1.4 | 312 | 3235 |  | 3791 E2E Stream receive |  |  |  | 3140 |  |  |  |  | 3791 E2E Stream receive | [9 empty] | 180005570
Paris BTE13 PIG flagship | 130/70 | 6.0.1.5 | 13 | 5136 |  | n/a |  |  |  | 4673 |  |  |  |  | n/a | [9 empty] | 171146846 |  | Radio is always on.
|  | 6.0.1.5 | 13 | 5268 |  | n/a |  |  |  | 4715 |  |  |  |  | n/a | [9 empty] | 171146849 |  | Radio is always on.
Paris RIC312 Pig flagship | 115/50 | 6.0.1.5 | 312 | 5612 |  | n/a |  |  |  | 5253 |  |  |  |  | n/a | [9 empty] | 180005570
Paris RIC312 Pig flagship | 115/50 | 6.0.1.5 | 312 | 5543 |  | n/a |  |  |  | 5124 |  |  |  |  | n/a | [9 empty] | 180005573
Milan RIC used for Z-power testing | 115/50 | 5.6.0.6 | 312 | 2242 | 2127 | 5446 |  | 5702 |  | 1514 |  |  |  |  | 5446 | 5702 | [8 empty] | 180309685
Milan RIC used for Z-power testing | 115/50 | 5.6.0.6 | 312 | 2261 | 2145 | 5474 |  | 5912 |  | 1536 |  |  |  |  | 5474 | 5912 | [8 empty] | 180309710
Zpower Napa Sapphire RIC312 63000-110 | 115/50 | 5.7.0.17 | 312 | 2370 | 2290 | 5490 |  | 6950 |  | 1630 |  |  |  |  | 5490 | 6950 | [8 empty] | 180309672
| 115/50 | 5.7.0.17 | 312 | 2360 | 2250 | 5540 |  | 7030 |  | 1590 |  |  |  |  | 5540 | 7030 | [8 empty] | 180309665
Quebec uRIC 312 | 115/50 | 7.0.0.6 | 312 | 3487 | 3357 | 4247 |  |  |  | 3292 |  |  |  |  | 4247 | [9 empty] | 180367412
Quebec uRIC 312 | 115/50 | 7.0.0.6 | 312 | 3665 | 3614 | 4492 |  |  |  | 3411 |  |  |  |  | 4492 | [9 empty] | 180367413
Quebec uRIC 312 | 115/50 | 7.0.0.8 | 312 | 3780 | 3740 | 4750 |  |  |  | 3520 |  |  |  |  | 4750 | [9 empty] | 180367412
Quebec uRIC 312 | 115/50 | 7.0.0.8 | 312 | 3940 | 3920 | 4760 |  |  |  | 3710 |  |  |  |  | 4760 | [9 empty] | 180367413
Quebec uRIC 312 | 115/50 | 7.0.0.9 | 312 | 1830 | 1800 | 2840 |  |  |  | 1620 |  |  |  |  | 2840 | [9 empty] | 180367412
Quebec uRIC 312 | 115/50 | 7.0.0.9 | 312 | 1810 | 1790 | 2820 |  |  |  | 1620 |  |  |  |  | 2820 | [9 empty] | 180367413
Quebec uRIC 312 | 115/50 | 7.0.1.4 | 312 | 2170 | 2180 | 3080 |  |  |  | 1940 |  | 2730 |  |  |  |  | 1920 | [7 empty] | 181416810 |  | Right - iPhone SE
Quebec uRIC 312 | 115/50 | 7.0.1.4 | 312 | 2110 | 2100 | 3030 |  |  |  | 1920 |  | 2680 |  |  |  |  | 1870 | [7 empty] | 181416809 |  | Left - iPhone SE
Quebec uRIC 312 | 115/50 | 7.0.2.3 | 312 | 1946 | 1939 | 2376 |  |  |  | 1720 |  | 2541 | [12 empty] | 182048504
Quebec uRIC 312 | 115/50 | 7.0.2.3 | 312 | 2006 | 1897 | 2888 |  |  |  | 1711 |  | 2550 | [12 empty] | 182048520
Rome RC BTE After PMIC drain | 130/70 | 7.1.0.5 | RC | 1867 | 1757 | 2902 | 2868 | n/a | n/a | 1628 | 2593 | 2574 |  | 1693 | [10 empty] | 181416625
Rome RC BTE After PMIC drain | 130/70 | 7.1.0.5 | RC | 1887 | 1706 | 2904 | 2882 |  |  | 1669 | 2613 | 2573 |  | 1894 | [10 empty] | 181416626
Rome RC BTE After PMIC drain | 130/70 | 7.1.2.3 | RC | 1775 | 1655 | 2773 |  |  |  | 1621 | 2437 | 2412 |  | 1553 | [10 empty] | 190645351
Rome RC BTE After PMIC drain | 130/70 | 7.1.2.3 | RC | 1735 | 1653 | 2763 |  |  |  | 1558 | 2404 | 2440 |  | 1558 | [10 empty] | 190645350
Sydney Custom ITE Dir Lt | 115/50 | 7.1.0.7 | RC | 1765 | 1610 | 2785 | 2841 | 1802 | 2775 | 1560 | 2450 | 2434 |  | 1543 | [10 empty] | 2911326913 |  | Sleep Current 0.413mA
Sydney Custom ITE Dir Rt | 115/50 | 7.1.0.7 | RC | 1812 | 1675 | 2928 | 2855 | 1639 | 2837 | 1582 | 2578 | 2544 |  | 1633 | [10 empty] | 2911326912 |  | Sleep Current 0.426mA
Quebec uRIC 312 | 115/50 | 7.0.4.1 | 312 | 2090 | 1940 | 3030 |  |  |  | 1940 |  | 2750 |  |  |  |  | 1880 | [7 empty] | 181416807
Quebec uRIC 312 | 115/50 | 7.0.4.1 | 312 | 2070 | 1970 | 3040 |  |  |  | 1950 |  | 2760 |  |  |  |  | 1860 | [7 empty] | 181416708
Quebec uRIC 312 | 115/50 | 7.1.2.5 | 312 | 1860 | 2100 | 2690 |  |  |  | 1900 | [6 empty] | 1880 |  |  |  | 1660 |  |  |  | 190673300 |  | Left - Pixel 3 Android | Still in phone connected mode
Quebec uRIC 312 | 115/50 | 7.1.2.5 | 312 | 1850 |  | 2780 |  |  |  | 1910 | [6 empty] | 1880 |  |  |  | 1640 |  |  |  | 182048573 |  | Right - Pixel 3 Android | Still in phone connected mode
Quebec uRIC 312 | 115/50 | 7.1.2.5 | 312 | 1860 | 1760 | 2790 |  |  |  | 1710 | [6 empty] | 1620 | [7 empty] | 190673300 |  | Left - iPhone SE
Quebec uRIC 312 | 115/50 | 7.1.2.5 | 312 | 1810 |  | 2730 |  |  |  | 1830 | [6 empty] | 1730 |  |  |  | 1670 |  |  |  | 182048573 |  | Right - iPhone SE | In fast adv mode after turn off BLE on iPhone
Sydney RIC R HR on not enabled | 115/50 | 6.8.0.4 | RC | 2226 | 2104 | 3080 | 3069 | 2150 | 4428 | 1846 | 2757 | 2761 |  | 2100 | 3080 | [9 empty] | 191323934 |  | Left - iPhone 5S
Sydney Custom ITE right - VBAT | 120/60 | 7.2.0.5 | RC | 625 | 578 | 916.6667 | 937.3333 | 590 | 973.3333 | 553.3333 | 844 | 848.3333 |  | 593.3333 |  |  | 543.3333 | [7 empty] | 2911327561 |  | right device iPhone X | Double Tap Off | 3.68 before PMIC =1.20
Sydney Custom ITE right - VHA Calc | 120/60 | 7.2.0.5 | RC | 1695 | 1554 | 2570 | 2632 | 1590 | 2740 | 1480 | 2352 | 2365 |  | 1600 | no NFMI |  | 1450 | [7 empty] | 2911327561 |  | Calc VHA current
Sydney Livio Edge Custom ITE right - VBAT | 120/60 | 7.2.0.7 | RC | 685 | 653 | 1020 | 998.3333 | 658 | 997.6667 | 606.6667 | 887 | 888 |  | 647.3333 |  |  | 606.6667 | [7 empty] | 2911327561 |  | right device iphone 5S 12.4.3 | Double Tap ON | 3.67 before PMIC -1.20 Kiwi fitting session disabled
Sydney Livio Edge Custom ITE right - VHA Calc | 120/60 | 7.2.0.7 | RC | 1875 | 1779 | 2880 | 2815 | 1794 | 2813 | 1640 | 2481 | 2484 |  | 1762 | no NFMI |  | 1640 | [7 empty] | 2911327561 |  | Calc VHA current
Sydney Livio Edge Custom ITE right - VBAT | 120/60 | 7.2.0.8 | RC | 687.3333 | 683 | 996.3333 | 1001 | 702 | 1005 | 664 | 905.6667 | 902 |  | 652 | [10 empty] | 2911327561 |  | right device iphone 5S 12.4.4 | Double Tap ON | 3.67 before PMIC -1.20 Kiwi fitting session disabled
Sydney Livio Edge Custom ITE right - VHA Calc | 120/60 | 7.2.0.8 | RC | 1882 | 1869 | 2809 | 2823 | 1926 | 2835 | 1812 | 2537 | 2526 |  | 1776 | [10 empty] | 2911327561 |  | Calc VHA current
Rome RC BTE After PMIC drain | 130/70 | 7.2.0.11 | RC | 1921 | 1885 | 2915 | 2864 | 1912 | 3017 | 1792 | 2649 | 2613 |  | 1762 | [12 empty] | Auto ON/OFF enabled, double tap enabled
Rome RC BTE After PMIC drain | 130/70 | 7.2.0.11 | RC | 1964 | 1825 | 2914 | 2878 | 1975 | 2931 | 1718 | 2692 | 2619 |  | 2174
Quebec uRIC | 115/50 | 7.4.0.4 | 312 | 2055 | 1902 | 3055 | 3050 |  |  | 1917 |  | 2780 |  |  |  |  | 1875 | [7 empty] | 200189554 |  | Auto ON/OFF enabled, double tap enabled
Quebec uRIC | 115/50 | 7.4.0.4 | 312 | 2055 | 1892 | 3028 | 3022 |  |  | 1907 |  | 2762 |  |  |  |  | 1858 | [7 empty] | 200189552 |  | Auto ON/OFF enabled, double tap enabled
SoundGear 2.4 Custom ITE DIR VHA | 90/30 | 7.4.0.9 | RC | 1860 | 1620 | 2760 |  |  |  | 1540 | [14 empty] | 2911329568 |  | wireless Kiwi session 2.02mA iPhone 7 (13.2)
SoundGear 2.4 Custom ITE DIR VHA | 90/30 | 7.4.1.1 | RC | 1760 | 1600 | 2670 |  |  |  | 1530 | [14 empty] | 2911329568 |  | wireless Kiwi session 2.03mA iPhone 7 (13.2)
Tulum Livio Edge AI 2400 RBTE VHA | 130/70 | 7.4.0.9 | RC | 1780 | 1620 | 2780 |  |  |  | 1560 | [14 empty] | 190645350 |  | iPhone 7 (13.2)
Tulum Livio Edge AI 2400 RBTE VHA | 130/70 | 7.4.0.9 | RC | 1790 | 1650 | 2800 |  |  |  | 1580 | [14 empty] | 190645351 |  | iPhone 7 (13.2)
Sydney Livio Edge Custom ITE right - VHA | 120/60 | 7.4.0.9 | RC | 1720 | 1585 | 2730 |  |  |  | 1540 | [14 empty] | 2911327561 |  | iPhone 7 (13.2)
SoundGear Phantom | 90/30 | 7.5.1.0 | RC | 1933 | 1839 | 2806 |  |  |  | 1813 | [14 empty] | 491132647 |  | iPhone 12 Pro (15.11) | 2021-12-21 | 3.62v PrePmic
SoundGear Phantom | 90/30 | 7.5.1.0 | RC | 2050 | 1855 | 2790 |  |  |  | 1830 | [14 empty] | 491132648
G23 2.4GHz ITC Custom Rechargeable | 130/70 | 8.2.2.1 | RC | 450-460 | 400 | 670 |  |  |  | 400 |  |  |  |  | no NFMI |  | 380-390 | [7 empty] | 2911334012 |  | Right device, iPhone 12 mini, Google Pixel 5 | 2022-05-19
G23 2.4GHz ITC Custom Rechargeable | 130/70 | 8.2.2.1 | RC | 460-470 | 400 | 680 |  |  |  | 400 |  |  |  |  | no NFMI |  | 380-390 | [7 empty] | 2911334013 |  | Left device, iPhone 12 mini, Google Pixel 5 | 2022-05-19
G23 2.4GHz ITC Custom Rechargeable | 130/70 | 8.2.2.1 | RC | 500-540 | 380-410 | 700-720 |  |  |  | 370-420 | [6 empty] | 360-390 | [9 empty] | BleTransmitPower = 2
|  |  |  | unstable | unstable
G23 2.4GHz ITC Custom Rechargeable | 130/70 | 8.2.2.1 | RC | 490-510 | 380-400 | 690-710 |  |  |  | 380-410 | [6 empty] | 360-380 | [9 empty] | BleTransmitPower = 2
G23 2.4GHz ITC Custom Rechargeable | 130/70 | 8.2.2.4 | RC | 580 | [12 empty] | 430 | [7 empty] | 2911334013
G23 2.4GHz ITC Custom Rechargeable Directional VBAT | 115/50 | 8.2.2.6 | 312 | 550-580 | 420-450 | 760-800 |  |  |  | 420-460 | [6 empty] | 400-430 |  |  | 177 |  |  |  |  | 2911334631
|  |  |  | 550-580 | 420-450 | 760-800 |  |  |  | 420-460 | [6 empty] | 400-430 |  |  | 173 |  |  |  |  | 2911334632
G23 2.4GHz ITC Custom Rechargeable Directional Telecoil VBAT | 115/50 | 8.2.2.6 | 312 | 540-580 | 420-460 | 770-820 |  |  |  | 420-460 | [6 empty] | 400-430 | [7 empty] | 2911334633
|  |  |  | 540-580 | 420-460 | 770-820 |  |  |  | 420-460 | [6 empty] | 400-430 | [7 empty] | 2911334634
Yellowstone Wireless 2.4Hz CIC Zinc Air | 115/50 | 8.3.0.9 | 312 | 1450-1550 | 1800-1900 | 2200-2400 |  |  |  | 1200-1300 | [6 empty] | 950-1000 |  |  | 483 SYS-15049 |  |  |  |  | 2911336716 (Right) |  |  | 2023-04-03
SoundGear Phantom | 90/30 | 7.10.0.7 | RC | 780 | 609 | 965 | 1034 | N/A | N/A | 580 | N/A | 850 | 895 | 598 | N/A | N/A | 590 | 648 | 906 | 0.012 |  |  |  |  | 2422210242 (right) |  | 2-way phone call with iPhone 7 (iOS 15.7) Samsung Galaxy A72 | 2023-06-19 | 3.65v Vbat (1.20vHA)
Note on K51: Bob Rusten: After 06's high Nominal now both aids stay near advertising drain.

## Genesis
Product | Matrix | Firmware Rev. | Battery Size | Advertising Current (µA) | 30ms Sync Current (µA) | iPhone Stream Current (µA) | iPhone Call Current (µA) | Nominal Current (µA) | E2E Transmit Current (µA) | E2E Receive Current (µA) | TVS Stream Current (µA) | MFA Stream Current (µA) | MFA Phone Call Current (µA) | BNR Stream Current (µA) | MNR + Stimulus Current (µA) | 90 dBSPL Stimulus Current (µA) | Slow Advertisement Current (µA) | Android connection Current (uA) | Android Streaming Current (uA) | Auto Sleep Current (uA) | Expected Days Typ Accessory | Actual Days Typ Accessory | Expected Days Non Accessory | Actual Days Non Accessory | Serial Numbers | Database ID Code | Notes | Date | NOTE: * = Calculated
Genesis Gummy Plus Flag | 115/50 | 8.2.0.6 | RC | 620 | 770 | 940 | 940 | 560 |  |  | 850 | [7 empty] | 560
Genesis Gummy Plus Micro RIC | 115/50 | 8.2.0.8 | RC | 494 | 437 | 795 |  | 430 | [10 empty] | 450 |  |  |  | 200195976 |  | iPhone 7 (14.3) Battery power 3.65v
Genesis Gummy Plus Micro RIC | 115/50 | 8.2.0.8 | RC | 490 | 435 | 790 |  | 425 | [13 empty] | 440 |  |  | 200195977 |  | iPhone 7 (14.3) Battery power 3.65v
Genesis 2.4GHz Micro RIC-R #2998 | 115/50 | 8.2.0.9.0 | RC | 420 | 380 | 730 |  | 380 | [8 empty] | 360 | 420 | 720 | [5 empty] | 200195977 |  | Left; iPhone 11 (14.4); Google Pixel 4 (Android 11)
|  |  |  | 420 | 370 | 720 |  | 370 | [8 empty] | 350 | 410 | 710 | [5 empty] | 200195976 |  | Right; iPhone 11 (14.4); Google Pixel 4 (Android 11)
Genesis 2.4GHz Micro RIC | 115/50 | 8.2.0.10.0 | RC | 460 | 410 | 800 |  | 420 | [8 empty] | 400 | 450-480 | 790 | [5 empty] | Right: 200195980 |  | iPhone 12 mini (14.4); Google Pixel 4a (Android 11) | 2021-03-17
|  |  |  | 460 | 410 | 760 |  | 410 | [8 empty] | 390 | 450-470 | 750 | [5 empty] | Left: 200195979
Genesis Gobstopper Plus Flag | 115/50 | 8.2.0.6 | RC | 750 | 931 | 1110 | 1060 | 690 | 750 | 960 | 990 | [5 empty] | 690
Genesis Macro RIC | 115/50 | 8.2.0.8.0 | RC | 530 | 470 | 835 |  | 470 | [8 empty] | 460 | [7 empty] | 190676066 |  | iPhone 12 mini (14.3)
190676067
Genesis Macro RIC | 115/50 | 8.2.0.8.0 | RC | 810 | 760 | 1150 |  | 760 | [8 empty] | 790
Genesis Macro RIC | 115/50 | 8.2.0.9.0 | RC | 760 | 710 | 1070 |  | 720 | [8 empty] | 700 | 520 | 820 | [5 empty] | 190676066 |  | Left; iPhone 11 (14.4); Google Pixel 4a (Android 11) | 2021-02-09
|  |  |  | 530 | 470 | 830 |  | 470 | [8 empty] | 460 | 520 | 760 | [7 empty] | Right; iPhone 11 (14.4); Google Pixel 4a (Android 11)
Genesis Macro RICs | 115/50 | 8.2.0.10.0 | RC | 540 | 480-540 | 850 |  | 470-530 | [8 empty] | 500 | 520-580 | 840 | [5 empty] | Right: 190676066 |  | iPhone 12 mini (iOS 14.4); Google Pixel 4a (Android 11) | 2021-03-17
|  |  |  | 540 | 480-550 | 870 |  | 480-540 | [8 empty] | 540 | 520-580 | 800 | [5 empty] | Left: 201527253
Genesis DVT2 Macro RIC | 114/50 | 8.2.0.21.0 | RC | 550 | 540-560 | 860-880 |  | 500-520 | [8 empty] | 490 | 550-560 | 850-860 | [5 empty] | 211758998 |  | iPhone 12 mini (iOS 14.7.1); Google Pixel 4a (Android 11) | 2021-09-17
Genesis DVT2 Micro RIC | 114/50 | 8.2.0.21.0 | RC | 500 | 450-490 | 810 |  | 460-480 | [8 empty] | 430 | 490-510 | 790-810 | [5 empty] | 210654047 |  | iPhone 12 mini (iOS 14.7.1); Google Pixel 4a (Android 11) | 2021-09-14
Genesis DVT4 Macro RIC | 115/50 | 8.2.0.29.0 | RC | 580-620 | 520-550 | 810-820 |  | 500-520 | [8 empty] | 490-510 |  |  | 176 |  |  |  |  | (Left) 221001940 |  | iPhone 12 mini (iOS …); Google Pixel 4a | 2022-04-19
Genesis DVT4 Macro RIC | 115/50 | 8.2.0.29.0 | RC | 580-620 | 520-550 | 820-830 |  | 520-550 | [8 empty] | 500-520 |  |  | 184 |  |  |  |  | (Right) 221001946 |  | iPhone 12 mini (iOS …); Google Pixel 4a | 2022-04-19
G23 DVT4 Micro RIC | 115/50 | 8.2.2.6 | RC | 490 | 440 | 730 |  | 440 | 470 | 470 | [6 empty] | 400 |  |  | 185 |  |  |  |  | 221254584 |  | iPhone 12 mini | 2022-08-15
G23 DVT5 Macro RIC | 115/50 | 8.2.2.9 | RC | 570-610 | 520-550 | 800-810 |  | 520-550 | [8 empty] | 470-500 |  |  | 182
|  |  |  | 580-600 | 520-550 | 810-820 |  | 520-550 | [8 empty] | 470-500 |  |  | 176 |  |  |  |  | 222102428 (Right)
222102427 (Left)
G23 DVT5 Macro RIC | 115/50 | 8.2.2.9 | RC | 580-610 | 520-540 | 790-810 |  | 520-540 | [8 empty] | 500-520 |  |  | 167
|  |  |  | 580-610 | 520-540 | 800-820 |  | 520-540 | [8 empty] | 500-520 |  |  | 173 |  |  |  |  | 222102429 (Right)
222102430 (Left)
G23 DVT5 Micro RIC | 115/50 | 8.2.2.9 | RC | 450-500 | 440-450 | 720-740 |  | 420-390 |  |  | 720-740 | 720-740 |  |  |  |  | 440-420 |  |  | 170
|  |  |  | 450-500 | 430-450 | 720-730 |  | 410-390 |  |  | 720-730 | 720-730 |  |  |  |  | 440-420 |  |  | 166 |  |  |  |  | (L) Wired: 222102849 , (R)Wireless: 222102851 |  | iPhone 14 iOS
G23 DVT5 Micro RIC (Double Tap enabled) | 115/50 | 8.2.2.11 | RC | 490-510 | 420-440 | 720-730 |  | 430-460 | [8 empty] | 420-440 |  |  | 215 |  |  |  |  | 222102848 |  | iPhone 12 mini | 2022-12-28
G23 DVT5 Micro RIC (DT disabled) | 115/50 | 8.2.2.11 | RC | 470-490 | 400-430 | 690-710 |  | 410-430 | [8 empty] | 390-410 |  |  | 217 |  |  |  |  | 222102848 |  | iPhone 12 mini | 2022-12-28
G23 DVT5 Macro RIC (DT enabled) | 115/50 | 8.2.2.11 | RC | 570-590 | 500-530 | 790-800 |  | 500-530 | [8 empty] | 490-510 |  |  | 224 |  |  |  |  | 222102429
G23 DVT5 Macro RIC (DT disabled) | 115/50 | 8.2.2.11 | RC | 530-550 | 470-500 | 750-770 |  | 460-490 | [8 empty] | 450-480 |  |  | 224 |  |  |  |  | 222102429
Yellowstone Yoohoo Flagship | 115/50 | 8.3.0.8.0 | Zinc air 312 | 2470-3240 | 1490-1710 | 2640-2850 | 2680-2810 | 1450-1510 | 2870 | 2750 | 2440 | 2470 |  | 2860 | 2901 | 2912 | 1440-1550 | 1660-1922 | 2680-2820 | PD-5856 |  |  |  |  | 230229767 |  | iPhone Xs and OnePlus 9 5G Config IP7 (TC in IP6) | 2023-02-07
4200
Yellowstone Yoohoo Flagship | 115/50 | 8.3.0.8.0 | Zinc air 312 | 2500-3210 | 151-168 | 2670-2780 | 2668-2789 | 1420-1530 | 2890 | 2770 | 2440 | 2440 |  | 2880 | 2800 | 2790 | 1450-1560 | 1590-1870 | 2750-2930 | PD-5856 |  |  |  |  | 230229766 |  | iPhone Xs and OnePlus 9 5G ConfigIP7 (TC in IP6) | 2023-02-07
4200
Yellowstone Yoohoo Flagship | 120/60 | 8.3.0.8.0 | Zinc air 312 | PD-5748
Yellowstone Yoohoo Flagship | 120/60 | 8.3.0.8.0 | Zinc air 312 | PD-5748
Yellowstone Yoohoo Flagship | 130/70 | 8.3.0.8.0 | Zinc air 312 | PD-5748
Yellowstone Yoohoo Flagship | 130/70 | 8.3.0.8.0 | Zinc air 312 | PD-5748
G 23 DVT 5 Macro RIC R with Yellowstone fw (DT enabled) | 115/50 | 8.3.0.9 | RC | 550-600 | 490-500 | 760-790 |  | 450-490 | 900-930 | 840-860 | [6 empty] | 450-500 | [7 empty] | 222102428
Yellowstone Yoohoo Flagship | 107/40 | 8.3.0.9 | Zinc air 312 | 1670 | 1260 | 2310 | 2910 | 1270 | 2650 | 2510 | 2240 | 2230 |  | 3150 | 1690 | 1280 | 1230 |  | 2420 | [5 empty] | 230229753 |  | current typically fluctuates ~50uA for each measurement | 2023-03-20
Yellowstone Yoohoo Flagship | 107/40 | 8.3.0.9 | Zinc air 312 | 1790 | 1340 | 2310 | 2170 | 1490 | 2620 | 2430 | 2130 | 2140 | 2120 | 3280 | 1700 | 1320 | 1260 |  | 2480 | [5 empty] | 230229755 |  | current typically fluctuates ~50uA for each measurement | 2023-03-20
Yellowstone Yoohoo Flagship | 115/50 | 8.3.0.9 | Zinc air 312 | 1720 | 1390 | 2280 | 2920 | 1340 | 2720 | 2650 | 2200 | 2240 | 2260 | 3320 | 1640 | 1890 | 1310 |  | 2400 | [5 empty] | 230229753 |  | current typically fluctuates ~50uA for each measurement | 2023-03-20
Yellowstone Yoohoo Flagship | 115/50 | 8.3.0.9 | Zinc air 312 | 1810 | 1430 | 2250 | 2230 | 1540 | 2650 | 2790 | 2090 | 2170 | 2210 | 3300 | 1850 | 3160 | 1290 |  | 2450 | [5 empty] | 230229755 |  | current typically fluctuates ~50uA for each measurement | 2023-03-20
Yellowstone Yoohoo Flagship | 120/60 | 8.3.0.9 | Zinc air 312 | 1770 | 1380 | 2270 | 2240 | 1380 | 2700 | 2590 | 2190 | 2250 |  | 2840 | 1260 | 4030 | 1190 |  | 2430 | [5 empty] | 230229753 |  | current typically fluctuates ~50uA for each measurement | 2023-03-20
Yellowstone Yoohoo Flagship | 120/60 | 8.3.0.9 | Zinc air 312 | 1810 | 1420 | 2240 | 2180 | 1520 | 2680 | 2580 | 2190 | 2250 |  | 3310 | 1750 | 4610 | 1250 |  | 2400 | [5 empty] | 230229755 |  | current typically fluctuates ~50uA for each measurement | 2023-03-20
YST DVT5 Macro RIC-R | 115/50 | 8.3.0.11 | RC | 500-520 |  |  |  |  | 840-860 | 760-780 |  |  |  | 920-930 |  |  | 410-430 | [7 empty] | 222102429 |  | 4.0V, No phone connection, no IMU enabled, just getting BNR and E2E stream currents | 2023-05-04
Note on I1: Sydney Robillard - Contractor: This measurement is taken after streaming from the iPhone with the aids still connected
Note on R1: Sydney Robillard - Contractor: This measurement is taken with the aids disconnected from any device
Note on N59: Abdi Negash: Old test case
Note on U59: Abdi Negash: Unable to test
Note on Q60: Abdi Negash: BNR activated possible Ignite connected
Note on N61: Abdi Negash: Old test case
Note on Q61: Abdi Negash: MNR
Note on U61: Abdi Negash: Unable to test
Note on Q62: Abdi Negash: BNR activated. Possible Ignite connected
Note on Q81: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: this high measurements are not fully accurate as BNR was enabled on the memory tested
Note on Q83: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: this high measurements are not fully accurate as BNR was enabled on the memory tested

## Wired
Product | Matrix | Firmware Rev. | Battery Size | Nominal Current (µA) | Expected Days Non Accessory | Actual Days Non Accessory | Serial Numbers | Database ID Code | Notes | NOTE: * = Calculated
London Wired RIC 312 | 115/50 | 4.4.1.1 | 312 | 1400 |  | *8.2 | 15373236 / 15373223
| 115/50 | 4.4.2.1 | 312 | 1400 |  | *8.2 | 15373236 / 15373223
| 123/60 |  | 312
| 130/70 |  | 312
London Wired ITC 312 | 115/50 |  | 312
| 120/60 | 4.4.1.1 | 312 | 1175 |  | *9.9 | 2911323612 / 2911323613
| 130/70 | 4.1.5.3 | 312 | 1275 | 7 | *9.0 | 2911323256 / 2911323257
London Wired ITE 13 | 115/50 |  | 13
| 123/60 |  | 13
| 130/70 | 4.1.5.3 | 13 | 1275 | 7 | *9.0 | 2911323254 / 2911323255
| 130/70 | 4.4.1.1 | 13 | 1225 | 7 | *9.4 | 2911323254 / 2911323255
London Wired CIC 312 | 115/50 | 4.1.5.3 | 10 | 1075 |  | *5.2 | 2911323621 / 2911323620
| 115/50 | 4.4.2.3 | 10 | 1050 | 7 | *5.3 | 2911323291 / 2911323292
| 123/60 |  | 10
| 130/70 |  | 10
London Wired IIC 10 | 110/40 | 4.1.5.3 | 10 | 1050 |  | *5.3 | 2911323614 / 2911323615
| 115/50 | 4.4.1.1 | 10 | 1075 |  | *5.2 | 2911323616 / 2911323617
| 123/60 |  | 10
| 130/70 |  | 10
Lyon Wired BTE | 110/40 |  | 13
| 115/50 |  | 13
| 123/60 |  | 13
| 130/70 | 4.11.0.1 | 13 | 1250 |  |  | 160184307
G23 Wired CIC | 115/50 | 8.2.2.6 | 312 | 755 |  |  | 2911334572 |  | Mfg Suite 26.0.3029, Pro Fit 1.0.30043.0, 1.2V Keithley supply
|  |  |  | 743 |  |  | 2911334573
G23 Wired IIC | 115/50 | 8.2.2.6 | 10 | 725 |  |  | 2911334574 |  | Mfg Suite 26.0.3029, Pro Fit 1.0.30043.0, 1.2V Keithley supply
|  |  |  | 752 |  |  | 2911334575
G23 Wired IIC | 110/40 | 8.2.2.9 | 10A | 645 |  |  | 2911334559 |  | Mfg Suite 26.0.3165 Pro Fit 1.0.30181.0 12v Keithley supply
G23 Wired IIC | 115/50 | 8.2.2.9 | 10A | 655 |  |  | 2911335065 |  | Mfg Suite 26.0.3165 Pro Fit 1.0.30181.0 12v Keithley supply
G23 Wired IIC | 123/60 | 8.2.2.10 | 10A | 647 |  |  | 2911333331 |  | Mfg Suite 26.0.5051 Pro Fit 1.0.30181.0, 1.2v Keithley supply
G23 Wired IIC | 115/50 | 8.2.2.10 | 10A | 660 |  |  | 2911335065 |  | Mfg Suite 26.0.5051 Pro Fit 1.0.30181.0 1.2v Keithley supply
G23 Wired IIC | 110/40 | 8.2.2.10 | 10A | 652 |  |  | 2911334559 |  | Mfg Suite 26.0.5051 Pro Fit 1.0.30181.0, 1.2v Keithley supply
G23 Wired IIC | 120/60 | 8.2.2.12 | 10A | 680 |  |  | 1823190266 |  | Mfg Suite (G23 VA Release) 26.0.6147.0, Pro Fit 1.0.35068.0 1.2V Keithley supply, office noise, sitting idle on a desk
|  |  |  | 675 |  |  | 1823190265
G23 Wired CIC | 110/40 | 8.2.2.12 | 10A | 673 |  |  | 1723191046 |  | Mfg Suite (G23 VA Release) 26.0.6147.0, Pro Fit 1.0.35068.0 1.2V Keithley supply, office noise, sitting idle on a desk
|  |  |  | 660 |  |  | 1723191045